## Sheet1
Technical Datasheet
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Major Surgical Instrument Set
1. General Description
1.1 The Major Surgical Instrument Set shall be suitable for general surgical procedures performed in hospitals and operating theatres.
1.2 The set shall include a complete range of instruments required for routine major surgical operations, particularly abdominal and soft tissue surgery.
1.3 All instruments shall be supplied new, unused, and ready for clinical use.
2. Material Requirements
2.1 All instruments shall be manufactured from medical-grade surgical stainless steel.
2.2 The stainless steel shall conform to AISI 410, AISI 420, AISI 440A/440B or equivalent surgical-grade standards.
2.3 Instruments shall be corrosion-resistant and suitable for repeated autoclave sterilization.
2.4 Instruments shall resist pitting, staining, oxidation, and degradation due to repeated cleaning and disinfection.
2.5 Instruments shall have a satin or matte anti-glare finish suitable for operating room conditions.
2.6 All instruments shall be free from defects such as burrs, cracks, sharp edges, and manufacturing imperfections.
2.7 All joints, hinges, ratchets, and locking mechanisms shall operate smoothly and maintain alignment.
2.8 Cutting instruments shall be capable of re-sharpening without loss of performance.
2.9 Each instrument shall be permanently marked with manufacturer identification and catalogue/reference number.
3. Set Components and Quantities
3.1 Forceps and Clamps
3.1.1 Dressing Forceps (14–16 cm) – Qty: 2
3.1.2 Tissue Forceps (14–18 cm) – Qty: 2
3.1.3 Artery Forceps (Straight, 14–20 cm) – Qty: 8
3.1.4 Artery Forceps (Curved, 14–20 cm) – Qty: 8
3.1.5 Kocher Forceps (14–18 cm) – Qty: 4
3.1.6 Allis Tissue Forceps (14–16 cm) – Qty: 4
3.1.7 Babcock Tissue Forceps (16 cm) – Qty: 2
3.1.8 Sponge Holding Forceps (18–24 cm) – Qty: 2
3.1.9 Backhaus Towel Forceps (9–13 cm) – Qty: 4
3.1.10 Rankin-Crile Hemostatic Forceps (16 cm Straight) – Qty: 2
3.1.11 Rankin-Crile Hemostatic Forceps (16 cm Curved) – Qty: 2
3.1.12 Rochester-Pean Hemostatic Forceps (16 cm Curved) – Qty: 2
3.1.13 Rochester-Ochsner Forceps (16 cm Straight) – Qty: 2
3.1.14 Halsted Mosquito Forceps (12.5 cm Straight) – Qty: 6
3.1.15 Halsted Mosquito Forceps (12.5 cm Curved) – Qty: 6
3.1.16 Dissecting Forceps (14.5 cm) – Qty: 2
3.1.17 Dissecting Forceps (20 cm) – Qty: 2
3.1.18 Tissue Forceps (20 cm, 1×2 Teeth) – Qty: 2
3.2 Needle Holders
3.2.1 Needle Holder (16–20 cm) – Qty: 2
3.2.2 TC Mayo-Hegar Needle Holder (18–20 cm) – Qty: 2
3.3 Scissors
3.3.1 Mayo Scissors (Straight, 14–18 cm) – Qty: 2
3.3.2 Mayo Scissors (Curved, 14–18 cm) – Qty: 2
3.3.3 Metzenbaum Scissors (14–18 cm) – Qty: 2
1 | 3.3.4 TC Wire Cutting Scissors (12 cm Angled) – Qty: 1
3.4 Scalpel Handles
3.4.1 Scalpel Handle No. 3 – Qty: 2
3.4.2 Scalpel Handle No. 4 – Qty: 2
3.5 Retractors
3.5.1 Richardson Retractor (Multiple sizes) – Qty: 3 sets
3.5.2 Deaver Retractor (300 × 50 mm) – Qty: 2
3.5.3 Volkmann Retractor (4-prong) – Qty: 2
3.5.4 Volkmann Retractor (6-prong) – Qty: 2
3.5.5 Cushing Vein Retractor – Qty: 2
3.5.6 US Army Double-Ended Retractor – Qty: 2 sets
3.5.7 Ribbon Retractor (25 mm) – Qty: 1
3.5.8 Ribbon Retractor (40 mm) – Qty: 1
3.5.9 Balfour Abdominal Retractor – Qty: 1
3.5.10 Langenbeck Retractors (Small) – Qty: 2
3.5.11 Langenbeck Retractors (Medium) – Qty: 2
3.6 Intestinal and Special Instruments
3.6.1 Doyen Intestinal Forceps – Qty: 4
3.6.2 Lane Tissue Forceps – Qty: 2
3.6.3 Grooved Director (14.5 cm) – Qty: 1
3.6.4 Probe with Eye (13 cm) – Qty: 1
3.6.5 Deschamps Ligature Needle (Right) – Qty: 1
3.6.6 Deschamps Ligature Needle (Left) – Qty: 1
3.7 Suction Instruments
3.7.1 Yankauer Suction Tube (28 cm) – Qty: 2
3.7.2 Frazier Suction Tube – Qty: 1
3.8 Trays and Accessories
3.8.1 Kidney Tray – Qty: 2
3.8.2 Perforated Stainless Steel Sterilization Instrument Tray with Lid – Qty: 1
3.9 Equivalent Instruments
3.9.1 Equivalent standard surgical instruments may be accepted provided they meet the same functional and quality requirements.
4. Sterilization Compatibility
4.1 All instruments shall be suitable for steam sterilization (autoclave).
4.2 Instruments shall withstand repeated sterilization cycles without deformation or performance loss.
4.3 Instruments shall be resistant to corrosion, chemicals, disinfectants, and cleaning agents.
5. Storage and Packing
5.1 Instruments shall be supplied in a stainless steel sterilization tray suitable for autoclaving.
5.2 Instruments shall be individually protected and securely packed to prevent damage during transport.
6. Performance Requirements
6.1 Instruments shall ensure smooth operation and precise alignment.
6.2 Ratchets and locking systems shall function securely.
6.3 Cutting instruments shall maintain sharpness under normal clinical use.
6.4 Instruments shall be suitable for routine major surgical procedures.
7. Warranty
7.1 The supplier shall provide a minimum one (1) year warranty against manufacturing defects.
8. Spare Parts and Replacement
8.1 Replacement instruments shall be available upon request.
8.2 The supplier shall ensure availability of spare parts and replacements for a reasonable period.
9. Quality Standards
9.1 ISO 13485 certification or equivalent shall be provided.
9.2 CE Mark and/or FDA approval is preferred.
9.3 Manufacturer shall provide material certification confirming stainless steel grade.
10. Technical Documentation
10.1 Product catalogue and full instrument list shall be provided.
10.2 Material specification including stainless steel grade shall be included.
10.3 Quality certificates (ISO, CE, FDA or equivalent) shall be provided.
10.4 Warranty statement shall be included.
10.5 Country of origin shall be clearly stated.
2 | MINOR SURGICAL SET
This specification defines a Standard Minor Surgical Instrument Set intended for use in primary healthcare facilities, outpatient clinics, emergency rooms, and dressing units.
The set is designed for routine minor surgical procedures including wound management, suturing, dressing, hemostasis, and basic tissue handling.
All instruments shall be reusable, manufactured from medical-grade stainless steel, and suitable for repeated autoclave sterilization.
2. General Technical Requirements
All instruments shall comply with the following minimum requirements:
Medical-grade stainless steel (AISI 420 or equivalent)
Compliance with ISO 7153-1
Autoclavable at 134°C
Satin or matt anti-glare finish
Corrosion-resistant and reusable
Free from manufacturing defects (burrs, cracks, sharp edges)
Ergonomic design suitable for clinical use
Reliable locking mechanisms for forceps and needle holders
Suitable for repeated sterilization cycles without performance degradation
3. Instrument Composition
3.1 Cutting Instruments
Scalpel Handle No. 3 — Qty: 1
Scalpel Handle No. 4 — Qty: 1
Mayo Scissors Straight (14–15 cm) — Qty: 1
Mayo Scissors Curved (14–16 cm) — Qty: 1
Suture Removal Scissors (Spencer type) — Qty: 1
3.2 Hemostatic Instruments
Mosquito Forceps Straight (12–13 cm) — Qty: 2
Mosquito Forceps Curved (12–13 cm) — Qty: 2
Artery Forceps (Pean) Straight (14–16 cm) — Qty: 2
Artery Forceps (Pean) Curved (14–16 cm) — Qty: 2
3.3 Grasping & Tissue Handling Instruments
Dressing Forceps — Qty: 1
Tissue Forceps (Toothed 1×2) — Qty: 1
Dissecting Forceps (Non-toothed) — Qty: 1
Sponge Holding Forceps (Foerster) — Qty: 1
Needle Holder (Mayo-Hegar) (14–16 cm) — Qty: 1
3.4 Exploration & Support Instruments
Surgical Probe / Director — Qty: 1
Backhaus Towel Clamp — Qty: 2
3.5 Dressing Accessories & Containers
Stainless Steel Kidney Dish (500–700 ml) — Qty: 1
Stainless Steel Gallipot (120–200 ml) — Qty: 1
Stainless Steel Instrument Tray (approx. 25 × 20 × 5 cm) — Qty: 1
Autoclavable Sterilization Box with Lid — Qty: 1
4. Packaging Requirements
Instruments supplied in a rigid autoclavable tray or sterilization container
Instruments properly arranged to prevent movement or damage
Clearly labeled contents list included
Batch/lot number for traceability required
Supplied non-sterile but ready for sterilization
Protected against moisture, corrosion, and mechanical damage during transport
5. Quality Standards
The supplier shall provide valid documentation including:
ISO 13485 Certificate (Medical Devices Quality Management System)
ISO 9001 Certificate (preferred)
ISO 7153-1 compliance statement
CE Mark (Medical Device Regulation compliant)
Declaration of Conformity
Certificate of Origin issued by competent authority
All certificates must be issued by recognized certification bodies.
6. Documentation Requirements
The supplier shall submit the following:
Manufacturer catalogue / technical datasheets
Packing list
Warranty certificate (minimum 24 months)
Sterilization compatibility statement (134°C autoclave)
User and maintenance instructions
Delivery note
Inspection and quality control report
7. Inspection and Acceptance Criteria
Final acceptance shall be based on:
Visual inspection of workmanship and finish quality
Verification of quantities and compliance with specification
Functional testing of instruments:
Cutting performance (scissors)
Locking mechanism integrity (forceps/needle holder)
Alignment and grip performance
Verification of packaging integrity
Review of all required documents and certifications
8. Warranty Requirements
Minimum warranty period: 24 months from acceptance
Covers manufacturing defects and material failures
Replacement of defective items at supplier cost
Full after-sales support during warranty period
9. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted, provided they fully comply with the technical requirements and are approved by the procuring authority.
3 | Orthopedic Surgical Instrument Set
1. General Description
This specification defines a Complete Orthopedic Trauma Instrument Set for use in operating rooms and orthopedic departments for fracture management, bone fixation assistance, and trauma surgery procedures.
2. Material Requirements
Medical-grade stainless steel (AISI 420 / AISI 410 or equivalent)
Autoclavable at 134°C
Corrosion resistant and reusable
Satin/matt anti-glare finish
High tensile and torsional strength suitable for orthopedic use
Free from defects, cracks, and sharp edges
3. Instrument Composition with Exact Dimensions
3.1 Bone Holding & Reduction Instruments
Bone Holding Forceps (Reposition type, serrated jaws) — 18–24 cm — Qty: 1
Reduction Forceps (Pointed jaw, self-centering type) — 18–22 cm — Qty: 1
Plate Holding Forceps (Dynamic compression type) — 18–20 cm — Qty: 1
Bone Hook (Single sharp hook, ergonomic handle) — 16–18 cm — Qty: 1
3.2 Bone Cutting & Shaping Instruments
Bone Rongeur Forceps (double action, curved jaw) — 16–20 cm — Qty: 1
Bone Cutter / Nibbler (heavy duty) — 18–22 cm — Qty: 1
Osteotomes (Flat, straight, hex handle):
6 mm × 15 cm
10 mm × 15 cm
12 mm × 15 cm
18 mm × 18 cm
Bone Curettes (sharp & blunt, double-ended):
Small: 14 cm
Medium: 16 cm
Large: 18 cm
3.3 Bone Preparation Instruments
Periosteal Elevator (Freer type, double ended) — 14–17 cm — Qty: 1
Bone Awl (cortical perforation, straight) — 14–16 cm — Qty: 1
Depth Gauge (AO standard type) — 14–16 cm — Qty: 1
3.4 Fixation & Implant Support Instruments
Screwdriver Set (AO compatible, hexagonal coupling):
2.5 mm
3.5 mm
4.5 mm
Length: 15–20 cm — Qty: 1 set
Drill Chuck with Key (manual, Jacobs type) — 14–16 cm — Qty: 1
Pin Cutter (K-wire cutter, heavy duty) — 18–20 cm — Qty: 1
3.5 Impact Instruments
Bone Mallet (stainless steel head, wooden or metal handle)
Head diameter: 40–50 mm
Total length: 25–30 cm — Qty: 1
Gigli Saw Handle Set with wire
Handle length: 12–14 cm — Qty: 1 set
3.6 Retracting Instruments
Langenbeck Retractor — 20 × 25 mm blade, length 18 cm — Qty: 1
Hohmann Retractor (various sizes):
Small: 14 cm
Medium: 16 cm
Large: 18 cm
Sharp Retractor — 16–18 cm — Qty: 1 set
3.7 Auxiliary Instruments
Instrument Tray / Sterilization Basket
Size: 40 × 20 × 6 cm (standard orthopedic tray) — Qty: 1
Silicone/foam instrument fixation insert — included
4. Performance Requirements
Precise mechanical alignment of all moving parts
High resistance to bending and torsional stress
Cutting edges capable of clean bone and wire cutting
Secure locking mechanisms for forceps
Stable grip under heavy orthopedic load
Suitable for repeated trauma procedures
Long service life under intensive sterilization cycles
5. Sterilization Requirements
Compatible with autoclave sterilization at 134°C
Resistant to repeated chemical and thermal disinfection
No corrosion or performance degradation after repeated cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed positions with labeling system
Tray includes protective lid and locking system
Each set includes full contents list
Batch/lot number traceability required
Supplied non-sterile but ready for sterilization
Moisture and impact resistant packaging
7. Quality Standards
Supplier shall provide:
ISO 13485 certificate
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (EU MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue with full technical specifications
Complete instrument list with dimensions
Material certification report
Warranty certificate
Sterilization compatibility statement (134°C)
User and maintenance manual
Packing list and QC inspection report
9. Inspection & Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance (rongeur, cutter, scissors)
Locking mechanism strength
Jaw alignment accuracy
Verification of full set completeness
Packaging integrity check
Compliance with all technical specifications
10. Warranty
Minimum warranty: 12 months from acceptance
Covers manufacturing defects and material failure
Free replacement of defective items
Availability of spare/replacement instruments
11. Compliance Statement
Equivalent products from internationally recognized manufacturers may be accepted, provided they fully comply with this specification and are approved by the procuring authority.
4 | OBSTETRICS & GYNECO (OBS/GYN) SURGICAL SET
1. General Description
This specification defines a Complete Obstetrics & Gynecology Surgical Instrument Set intended for use in:
Maternity wards
Delivery rooms
Operating theatres
Emergency obstetric care units
The set supports procedures including:
Vaginal examination
Assisted delivery procedures
Episiotomy repair
Uterine exploration
Dilatation & curettage (D&C)
Minor gynecological surgical procedures
Emergency obstetric interventions
2. Material Requirements
Medical-grade stainless steel (AISI 420 / AISI 304 or equivalent)
Corrosion resistant and autoclavable at 134°C
Satin / matt anti-glare finish
Smooth surfaces free from burrs or defects
High mechanical strength for repeated surgical use
Durable joints, ratchets, and locking mechanisms
3. INSTRUMENT COMPOSITION (FULL SET)
3.1 Vaginal Examination & Exposure Instruments
Sims Vaginal Speculum, medium (80 × 30 mm blade approx.) — Qty: 1
Cusco Vaginal Speculum, medium size — Qty: 1
Auvard Weighted Vaginal Speculum — Length 25–28 cm — Qty: 1
Breisky Vaginal Retractor — Length 18–20 cm — Qty: 1
3.2 Grasping & Cervical Instruments
Tenaculum Forceps (single hook, 20–22 cm) — Qty: 1
Allis Tissue Forceps (4×5 teeth, 15 cm) — Qty: 2
Ovum Forceps (24–26 cm, fenestrated jaws) — Qty: 2
Sponge Holding Forceps (Foerster, straight/curved, 24–26 cm) — Qty: 2
Ring Forceps (Uterine Dressing Forceps, 24–26 cm) — Qty: 1
3.3 Hemostatic Instruments
Kelly Hemostatic Forceps, straight (14–16 cm) — Qty: 2
Kelly Hemostatic Forceps, curved (14–16 cm) — Qty: 2
Crile Forceps (14–16 cm) — Qty: 2
3.4 Needle Holding & Suturing Instruments
Mayo-Hegar Needle Holder (16–20 cm) — Qty: 2
3.5 Cutting Instruments
Mayo Scissors Curved (14–17 cm) — Qty: 1
Episiotomy Scissors Curved (14–16 cm) — Qty: 1
Umbilical / Cord Scissors (14–16 cm) — Qty: 1
3.6 Uterine Instruments (D&C Set)
Uterine Curette Sharp (sizes 2–10 mm) — Qty: 1 set
Uterine Curette Blunt (sizes 2–10 mm) — Qty: 1 set
Sims Uterine Sound (25–30 cm) — Qty: 1
Hegar Uterine Dilators (1–10 mm graduated set) — Qty: 1 set
3.7 Retractors
Sims Vaginal Retractor — 18–20 cm — Qty: 1
Richardson Retractor — 20–22 cm — Qty: 1
Lateral Vaginal Retractor — 18–20 cm — Qty: 1
3.8 Auxiliary Instruments & Containers
Kidney Dish, stainless steel — Qty: 2
Gallipot, stainless steel (120–200 ml) — Qty: 4
Instrument Tray, stainless steel (40 × 20 × 6 cm) — Qty: 1
Sterilization Box with Lid (autoclavable) — Qty: 1
4. Performance Requirements
Smooth articulation and precise mechanical movement
Secure grip and alignment of jaws
High resistance to repeated sterilization cycles
Reliable cutting performance for scissors
Strong locking mechanisms for forceps
Suitable for obstetric emergencies and surgical procedures
5. Sterilization Requirements
Compatible with autoclave sterilization at 134°C
Resistant to corrosion and chemical disinfectants
No degradation after repeated sterilization cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed layout with identification
Tray includes protective lid and locking system
Full contents list included
Batch/lot traceability required
Supplied non-sterile but ready for sterilization
7. Quality Standards
Supplier shall provide:
ISO 13485 certification
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue with specifications
Full instrument list with dimensions
Material certificates
Warranty certificate
Sterilization compatibility statement
User & maintenance manual
Packing list & inspection report
9. Inspection and Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance
Locking mechanism integrity
Jaw alignment and grip
Verification of full set completeness
Packaging integrity check
Compliance with technical specification
10. Warranty
Minimum warranty: 24 months
Covers manufacturing defects and material failure
Replacement of defective instruments free of charge
After-sales support throughout warranty period
11. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted provided full compliance with this specification and approval of the procuring authority.
5 | ENT Surgical Instrument Set
1. General Description
This specification defines a Complete ENT (Ear, Nose & Throat) Surgical Instrument Set intended for use in:
Hospitals
ENT clinics
Operating rooms
Emergency departments
The set supports diagnostic and surgical ENT procedures including:
Ear examination and procedures
Nasal examination and nasal surgery
Throat and laryngeal procedures
Foreign body removal
Minor ENT surgical interventions
2. Material Requirements
All instruments shall be manufactured from:
Medical-grade stainless steel (AISI 420 / AISI 304 or equivalent)
Corrosion-resistant and suitable for repeated sterilization
Autoclavable at 134°C
Satin/matt anti-glare finish
Smooth surfaces free from defects, burrs, or sharp edges
High durability for repeated clinical use
3. COMPLETE INSTRUMENT COMPOSITION
3.1 Diagnostic ENT Instruments
Head Mirror (adjustable head band type) — Qty: 1
Tuning Fork Set (128 Hz, 256 Hz, 512 Hz) — Qty: 1 set
Tongue Depressor (stainless steel, adult type) — Qty: 2
Auris Otoscope Set (with ear specula) — Qty: 1 set
3.2 Ear Instruments
Ear Specula Set (2 mm – 8 mm assorted sizes) — Qty: 1 set
Ear Curettes (sharp & blunt, assorted sizes) — Qty: 1 set
Jobson Horne Probe with ring curette — Qty: 1
Alligator Ear Forceps (straight, 14–16 cm) — Qty: 1
Aural Forceps (bayonet type, 14–16 cm) — Qty: 1
Ear Syringe (metal, 200–300 ml) — Qty: 1
3.3 Nasal Instruments
Nasal Speculum (Thudicum type, small/medium/large) — Qty: 1 set
Killian Nasal Speculum — Length 14–16 cm — Qty: 1
Tilley Nasal Forceps — Length 15–18 cm — Qty: 1
Hartmann Nasal Dressing Forceps — Length 15–18 cm — Qty: 1
Nasal Septum Elevator — Length 14–16 cm — Qty: 1
Nasal Suction Cannula (Frazier type, 6–12 Fr) — Qty: 1 set
3.4 Throat & Oral Instruments
Boyle-Davis Mouth Gag (adult size) — Qty: 1
Tonsil Holding Forceps — Qty: 1
Magill Forceps (adult, 20–25 cm) — Qty: 1
Laryngeal Mirror Set (assorted sizes) — Qty: 1 set
Tongue Forceps — Qty: 1
3.5 Cutting & General Instruments
ENT Scissors (straight, 14–16 cm) — Qty: 1
Metzenbaum Scissors (curved, 14–18 cm) — Qty: 1
Scalpel Handle No. 3 — Qty: 1
Scalpel Handle No. 4 — Qty: 1
Artery Forceps Straight (14–16 cm) — Qty: 2
Artery Forceps Curved (14–16 cm) — Qty: 2
3.6 Grasping & Holding Instruments
Dressing Forceps — Qty: 1
Tissue Forceps (toothed 1×2) — Qty: 1
Needle Holder (Mayo-Hegar, 14–16 cm) — Qty: 1
Sponge Holding Forceps (Foerster, 20–24 cm) — Qty: 1
3.7 Suction & Auxiliary Instruments
Suction Tube Set (Frazier type, sizes 6–12 Fr) — Qty: 1 set
Kidney Dish (stainless steel) — Qty: 2
Gallipot (120–200 ml) — Qty: 2
ENT Instrument Tray (40 × 20 × 6 cm) — Qty: 1
Sterilization Box with Lid (autoclavable) — Qty: 1
4. Performance Requirements
Smooth mechanical performance of all instruments
Accurate alignment of jaws and tips
High cutting efficiency and durability
Secure grip and ergonomic handling
Suitable for ENT surgical and diagnostic procedures
Resistant to repeated sterilization cycles
5. Sterilization Requirements
Fully compatible with autoclave sterilization at 134°C
Resistant to corrosion and chemical disinfectants
No degradation after repeated sterilization cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed labeled positions
Tray includes lid and locking system
Full contents list included
Batch/lot traceability required
Supplied non-sterile but ready for sterilization
7. Quality Standards
Supplier shall provide:
ISO 13485 certificate
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue and technical datasheets
Full instrument list with specifications
Material certificates
Warranty certificate
Sterilization compatibility statement (134°C)
User & maintenance manual
Packing list and inspection report
9. Inspection and Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance
Jaw alignment
Locking mechanisms
Verification of full set completeness
Packaging integrity check
Compliance with technical specifications
10. Warranty
Minimum warranty: 24 months
Covers manufacturing defects and material failure
Free replacement of defective instruments
Full after-sales support during warranty period
11. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted provided they fully comply with this specification and are approved by the procuring authority.
6 | Defibrillator
Description of Functions
To be used in emergency & critical care departments to meets various resuscitation and monitoring needs.
Technical Specifications
Defibrillator with complete accessories operated with AC power supply and internal battery
Defibrillation function:
It shall be a manual defibrillator for external defibrillation
Able to perform synchronized defibrillation and non-invasive pacing therapy.
Defibrillation energy selection:
External biphasic: 50 - 270J
External Paediatric: 2 - 20J
System shall be user friendly, lightweight and easily transportable.
Waveform shape: biphasic.
The defibrillator paddles shall be easily interchangeable among adult, child, infant and internal paddles. It shall come with at least adult and paediatric paddles.
Can be used for neonatal/paediatric and adult defibrillation.
The unit shall be able to perform defibrillation and monitoring by using disposable electrodes.
Shall have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
Recharge time shall not be held longer than 15 seconds before discharge.
Energy charge & discharge and other selection/control buttons shall be available at the paddle handles.
ECG monitoring function:
Shall have a 3-leads ECG, Lead I, II & III, monitoring capability protected from defibrillation by mean of ECG electrodes and through-the-paddles monitoring
With heart rate display and alarms
With Lead-fault indicator
Shall have an integrated thermal printer/ recorder with paper speed of 25mm/sec
General function:
Shall have LCD that displaying at least dual ECG channel, HR, battery status, shock indicator and various data. Bidder to specify size of LCD screen and the no. of waveforms which can be displayed.
Shall have audio and visual alarms. (Please indicate in the next column type of alarms available)
Shall have HR limit and shockable rhythms alarms
Shall have a rechargeable battery when it is fully charged it shall deliver approximately 40 - 50 discharges or 2 hours of continuous ECG monitoring. Bidder to specify the type of battery used and number of discharge and monitoring hour.
Must have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
The unit shall be portable and come with a carrying bag able to keep all required accessories and consumables.
Please indicate the weight in kilogram (KG) of the unit included all accessories and carrying case. It shall be within 8kg
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
· Rechargeable battery, 1 piece on the unit
· Thermal paper x 2 rolls/sets
· Power cord x 1 set
· 3 wire ECG cable x 1 set for ECG monitoring
· Disposable ECG electrodes, 50 pieces
. Paddle for adult and paediatric
· Carry Bag/case x 1 set
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Environmental factors
Thu unit shall be capable of operating in ambient temperature of 20-50 deg C and relative humidity of 80%
Standards and safety requirements
Should be FDA or CE approved product certificate.
User Training
Must provide user operation and service training
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
Certificate of calibration and inspection from factory.
7 | Mobile Operating Lamp
1. General Description
Mobile LED operating lamp designed to provide high-intensity, shadow-free illumination for minor surgical procedures and clinical examinations.
2. Type
Mobile / portable floor-standing surgical light
Suitable for operating rooms, emergency rooms, and clinics
3. Light Source
LED technology
High illumination intensity with long lifespan
Shadow-free lighting system
4. Illumination
Adjustable light intensity
High brightness suitable for surgical procedures
Color temperature suitable for medical use (white daylight range)
5. Mobility and Structure
Mounted on stable mobile base with wheels
Easy movement and positioning
| Lockable wheels for stability during use |  |  |  |  | The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-90%
Adjustable arm for flexible positioning
6. Power Supply
AC power supply: 220–240V, 50/60 Hz
Low power consumption
Backup option (if available) via battery system (optional)
7. Control System
Manual brightness adjustment
On/off control panel
Easy-to-use interface for clinical environments
8. Safety and Performance
Low heat emission to ensure patient comfort
Stable illumination without flickering
Designed for continuous medical use
9. Environmental Requirements
Operating temperature: 10°C – 40°C
Humidity: up to 85% non-condensing
Suitable for standard hospital environments
10. Warranty
2 years full warranty covering manufacturing and electrical faults
11. Spare Parts
Spare parts availability for 10 years minimum
Includes LED modules, power supply units, and control components
12. Quality Standards
ISO 13485 (mandatory)
ISO 9001 preferred
CE Mark and/or FDA approval required or equivalent certification
Compliance with medical electrical safety standards
13. Technical Documents
Product catalogue and technical datasheet
Quality certificates (ISO, CE, FDA or equivalents)
Manufacturer authorization (if applicable)
Warranty statement
Spare parts availability confirmation
List of similar installations (if available)
8 | Portable Ultrasound Machine (1 Convex Probe, 2 Probe Sockets)
1. Equipment Description
1.1 A general-purpose, fully digital Black & White (B&W) Ultrasound Imaging System.
2. Operational Requirements
2.1 The system shall operate on AC mains power supply.
3. System Configuration
3.1 Main ultrasound unit.
3.2 One (1) ultrasound probe.
3.3 One (1) Black & White thermal printer.
3.4 One (1) Ultrasound gel warmer.
4. Technical Specifications
4.1 General Requirements
4.1.1 The system shall be based on the latest technology and incorporate an all-digital beam former.
4.1.2 The system shall be a standalone ultrasound machine with an integrated lightweight mobile cart.
4.1.3 The system shall support the following applications:
Obstetrics and Gynecology (OB/GYN).
Abdominal imaging.
Peripheral vascular imaging.
Small parts imaging.
4.2 Display and Monitor
4.2.1 The system shall be equipped with a flat LCD monitor of at least 12 inches.
4.2.2 The monitor shall have tilt and swivel functionality.
4.3 Imaging Modes
4.3.1 B-Mode.
4.3.2 M-Mode.
4.3.3 B/M Mode.
4.3.4 2B Mode.
4.3.5 2D Mode.
4.4 Probe Connectivity
4.4.1 The system shall have a minimum of two (2) active probe ports for convenient operation.
4.5 Image Quality and Processing
4.5.1 The system shall provide 256 gray scale levels for high contrast resolution.
4.5.2 The system shall provide controls for:
Imaging depth.
Gain adjustment.
Time Gain Compensation (TGC).
Body markers with transducer position indication.
4.5.3 The system shall have real-time continuous dynamic focusing.
4.5.4 The system shall provide image zoom functionality.
4.5.5 The system shall provide image freeze functionality.
4.5.6 The system shall support text annotation.
4.6 Measurements and Calculations
4.6.1 The system shall include a comprehensive software package for OB/GYN and general imaging calculations.
4.6.2 The system shall provide measurement and calculation capabilities for:
Distance.
Area.
Volume.
Heart rate.
Circumference.
4.7 Advanced Imaging Features
4.7.1 The system shall include Tissue Harmonic Imaging (THI).
4.7.2 Near and far gain adjustment shall be available.
4.7.3 Contrast adjustment shall be available.
4.7.4 Focus shall be automatically adjustable.
4.8 User Interface
4.8.1 The system shall have an alphanumeric keyboard.
4.8.2 The system shall provide easily accessible scan controls.
4.8.3 The system shall include a trackball.
4.8.4 The system shall include a status display.
4.9 Memory and Storage
4.9.1 Cine memory shall support a minimum of 250 frames for cine loop playback.
4.9.2 Frame rate shall be not less than 50 frames per second (fps).
4.9.3 Display depth shall be at least 28–30 cm.
4.9.4 Dynamic range shall be selectable up to approximately 165 dB.
4.9.5 Internal image storage capacity shall be at least 200 patient images.
4.9.6 The system shall include an integrated CD writer.
4.9.7 The system shall support DICOM connectivity.
4.9.8 The system shall provide future upgradeability.
4.10 Probe Requirements
4.10.1 One (1) Convex probe with a frequency range of 2–5 MHz shall be supplied.
4.10.2 The probe shall be suitable for:
OB/GYN applications.
Abdominal imaging applications.
5. Accessories, Spare Parts, and Consumables
5.1 Accessories
5.1.1 One (1) Black & White thermal printer with one roll of high-density recording paper.
5.1.2 One (1) Ultrasound gel warmer.
5.2 General Requirements
5.2.1 All standard accessories required for normal operation shall be included.
5.2.2 All consumables necessary for operation shall be included.
5.2.3 All standard spare parts, tools, cleaning materials, and lubrication materials required for operation and routine maintenance shall be included in the offer.
6. Operating Environment
6.1 The equipment shall be designed for normal storage and operation under environmental conditions prevailing in Sudan.
6.2 Power supply requirements:
220–240 VAC.
50 Hz.
6.3 The equipment shall be supplied with an appropriate power plug.
6.4 The power cable length shall be not less than 3 meters.
6.5 A suitable UPS shall be supplied with:
Voltage regulation.
Spike protection.
Minimum 30 minutes backup time.
7. Standards and Safety Requirements
7.1 The manufacturer shall provide a valid ISO 13485:2003/AC:2007 certificate.
7.2 The offered product shall be certified with at least one of the following:
CE Certificate (93/42/EEC Directive).
US FDA Approval.
TUV Certification.
7.3 The equipment shall comply with:
IEC 60601-2-37 – Medical Electrical Equipment – Particular Requirements for the Basic Safety and Essential Performance of Ultrasonic Medical Diagnostic and Monitoring Equipment.
8. User Training
8.1 The supplier shall provide user training covering equipment operation.
8.2 The supplier shall provide training on routine maintenance procedures.
9. Warranty
9.1 The supplier shall provide a comprehensive warranty covering all parts and labor for a period of two (2) years from the date of final acceptance.
10. Maintenance Services During Warranty Period
10.1 Preventive maintenance shall be provided throughout the warranty period.
10.2 Corrective and breakdown maintenance shall be provided whenever required during the warranty period.
11. Documentation
11.1 One (1) User (Operating) Manual in English.
11.2 One (1) Service (Technical/Maintenance) Manual in English.
11.3 A list of essential spare parts and accessories including:
Part numbers.
Unit prices/costing.
9 | Electrosurgical Diathermy Unit with Trolley
1. Description of Function
1.1 The Diathermy Machine (Electrosurgical Unit) shall be used for surgical cutting and for controlling bleeding by coagulation (haemostasis) at the surgical site.
2. Operational Requirements
2.1 The unit shall operate on AC power supply suitable for use in operating theatres.
3. System Configuration
3.1 One (1) Electrosurgical Unit (ESU/Diathermy Machine) with a maximum output power of 300 W.
3.2 Complete set of accessories required for operation.
4. Technical Specifications
4.1 General Requirements
4.1.1 The unit shall be equipped with a solid-state, microprocessor-controlled frequency generator.
4.1.2 The unit shall provide monopolar and bipolar outputs electrically isolated from ground.
4.1.3 The operating frequency shall be greater than 320 kHz.
4.2 Monopolar Modes
4.2.1 The unit shall provide the following monopolar operating modes:
Pure Cut.
Blend Cut.
Soft Coagulation.
Contact Coagulation.
Spray Coagulation.
4.2.2 Maximum monopolar cutting power output shall be 300 W.
4.2.3 Maximum monopolar coagulation power output shall be 100 W.
4.3 Bipolar Modes
4.3.1 The unit shall provide bipolar cut mode.
4.3.2 The unit shall provide bipolar coagulation mode.
4.3.3 Maximum bipolar power output shall be 100 W.
4.4 Switching Functions
4.4.1 The unit shall support hand-switch operation when button-activated electrodes are connected.
4.4.2 A foot switch shall be supplied and capable of operating both monopolar and bipolar modes.
4.4.3 Yellow buttons/pedals shall activate the cutting function.
4.4.4 Blue buttons/pedals shall activate the coagulation function.
4.5 Safety Features
4.5.1 The grounding pad (return electrode) shall be continuously monitored for proper patient connection.
4.5.2 The system shall provide automatic generator deactivation in the event of grounding pad/return electrode failure.
4.5.3 The unit shall be protected against defibrillator discharges.
4.6 Controls and Display
4.6.1 The front control panel shall allow:
Mode selection.
Power adjustment.
ON/OFF operation.
4.6.2 The display shall indicate:
Output power settings.
System errors.
Electrode connection/failure alarms.
4.6.3 The unit shall provide audible and visual activation indicators.
4.6.4 The unit shall provide audible and visual alarm indicators.
4.6.5 The unit shall include a self-test mode.
4.7 Trolley
4.7.1 The unit shall be supplied with a dedicated trolley designed to securely accommodate the generator.
4.7.2 The trolley shall include drawers and storage space for accessories.
5. Accessories, Spare Parts and Consumables
5.1 General Requirement
5.1.1 All standard accessories, consumables, spare parts and cables required for proper operation of the equipment shall be included in the offer.
5.2 Accessories
5.2.1 One (1) Foot Switch comprising:
Two pedals (Yellow for Cut and Blue for Coagulation).
Connecting cable.
5.2.2 Two (2) Reusable Grounding Pads/Return Electrodes complete with:
Adult and pediatric sizes.
3-meter connecting cable.
5.2.3 Two (2) Reusable Monopolar Electrode Handles, foot-switch controlled, complete with connecting cables.
5.2.4 Two (2) Reusable Monopolar Electrode Handles, finger-switch controlled, complete with connecting cables.
5.2.5 One (1) Set of reusable monopolar electrodes including:
Needle electrode.
Knife electrode.
Ball electrode.
Wire loop electrode.
5.2.6 Two (2) Reusable Bipolar Forceps, foot-switch controlled, complete with connecting cables, including:
Short straight type.
Tip-angled type.
5.2.7 Two (2) Reusable Bipolar Forceps, foot-switch controlled, complete with connecting cables, including:
Long straight type.
Tip-angled type.
5.2.8 One (1) Mobile trolley fitted with:
Four (4) antistatic swivel castors.
Two (2) castors with brakes.
Drawer for accessories.
Storage compartment for foot switch/pedal.
6. Operating Environment
6.1 Storage Conditions
6.1.1 The unit shall be capable of continuous storage under:
Temperature range: 10°C to 60°C.
Relative humidity: 30% to 90%.
6.2 Operating Conditions
6.2.1 The unit shall operate under:
Temperature range: 10°C to 40°C.
Relative humidity: 30% to 75%.
6.3 Power Supply
6.3.1 Input power supply shall be:
220–240 VAC.
50 Hz.
Single phase.
6.3.2 The unit shall be supplied with a Schuko-type power plug.
7. Standards and Safety Requirements
7.1 The manufacturer shall provide a valid ISO 13485:2003/AC:2007 certification for Medical Devices.
7.2 The offered product shall have at least one of the following certifications:
CE Marking (93/42/EEC Medical Device Directive).
US FDA Approval.
TUV Certification.
7.3 The equipment shall comply with:
IEC 60601-2-2: Medical Electrical Equipment – Part 2-2: Particular Requirements for the Safety of High-Frequency Surgical Equipment.
8. User Training
8.1 The supplier shall provide comprehensive user training for equipment operators.
8.2 Training shall cover:
All operational functions of the equipment.
Proper use of accessories.
Routine user checks.
Basic maintenance procedures expected from users.
9. Warranty
9.1 The supplier shall provide a comprehensive warranty covering all parts and labor for a minimum period of two (2) years from the date of acceptance.
10. Maintenance Service During Warranty Period
10.1 The supplier shall provide Planned Preventive Maintenance (PPM) throughout the warranty period.
10.2 Corrective and breakdown maintenance services shall be provided whenever required during the warranty period.
11. Documentation
11.1 Two (2) sets of User (Operating) Manuals in English:
One hard copy.
One soft copy.
11.2 Two (2) sets of Service (Technical/Maintenance) Manuals in English:
One hard copy.
One soft copy.
11.3 A detailed list of important spare parts, accessories and consumables shall be provided, including:
Part numbers.
Unit prices/costing.
10 | Hydraulic Surgical Operating Table (Non-Electrical)
Technical Specifications for Hydraulic Operating Table
1. Description of Function
1.1 The Hydraulic Operating Table shall be designed for performing surgical procedures without the need for electrical power.
1.2 The table shall utilize a hydraulic mechanism for positioning and adjustment functions.
2. Operational Requirements
2.1 The operating table shall be suitable for general surgical procedures.
2.2 The table top shall be X-ray translucent and suitable for fluoroscopic procedures.
3. System Configuration
3.1 One (1) Hydraulic Operating Table complete with all standard accessories.
4. Technical Specifications
4.1 General Construction
4.1.1 The table shall have a four-section tabletop with a divided foot section.
4.1.2 The table shall be mounted on castors for mobility.
4.1.3 The table shall be equipped with an efficient braking system to ensure stability during surgery.
4.1.4 The tabletop shall be constructed from high-pressure laminate material suitable for:
X-ray penetration.
Fluoroscopy procedures.
4.1.5 The mattress shall be radiolucent and suitable for fluoroscopic imaging.
4.1.6 The frame casing and central support column shall be constructed from hygienic stainless steel.
4.2 Hydraulic Functions
4.2.1 The following movements shall be hydraulically operated:
Height adjustment.
Back section adjustment.
Lateral tilt.
Trendelenburg position.
Reverse Trendelenburg (Anti-Trendelenburg) position.
4.2.2 Foot section and head section adjustments may be manually operated.
4.2.3 The table shall be equipped with a manual position selector.
4.2.4 The position selector shall be interchangeable between the head and foot ends of the table.
4.3 Dimensions and Adjustments
(All dimensions are approximate with ±10% tolerance.)
4.3.1 Height adjustment range:
730 mm to 1040 mm.
4.3.2 Lateral tilt:
±15° minimum.
4.3.3 Back section adjustment:
From -15° to +70°.
4.3.4 Foot section adjustment:
From -90° to 0°.
Detachable.
4.3.5 Trendelenburg position:
Minimum 25°.
4.3.6 Reverse Trendelenburg (Anti-Trendelenburg):
Minimum 25°.
4.3.7 Head section adjustment:
From -40° to +30°.
Detachable.
4.3.8 Maximum table width:
Approximately 555 mm.
4.3.9 Overall table length:
Approximately 1950 mm.
4.4 Patient Capacity
4.4.1 The table shall support a maximum patient weight of approximately 200 kg.
5. Accessories, Spare Parts and Consumables
5.1 Accessories
5.1.1 One (1) Pair of padded arm rests complete with straps and clamps.
5.1.2 One (1) Anesthesia screen complete with clamps.
5.1.3 One (1) Pair of side supports complete with clamps.
5.1.4 One (1) Pair of knee crutches complete with clamps.
5.1.5 One (1) X-ray cassette tray.
5.1.6 One (1) Kidney bridge.
5.1.7 One (1) Stainless steel bowl complete with clamps.
5.1.8 One (1) Infusion rod (IV pole) complete with clamp.
5.1.9 One (1) Pair of leg supports.
5.2 General Requirements
5.2.1 All standard accessories required for normal operation shall be included.
5.2.2 All consumables, spare parts and components necessary for proper operation shall be included in the offer.
5.2.3 All standard tools, cleaning materials and lubrication materials required for operation and routine maintenance shall be included.
6. Operating Environment
6.1 The offered product shall be designed for storage and operation under environmental conditions prevailing in Sudan.
6.2 The equipment shall be capable of normal operation under local climatic conditions including temperature and humidity variations.
7. Standards and Safety Requirements
7.1 The manufacturer shall provide a valid ISO 13485:2003/AC:2007 certification for Medical Devices.
7.2 The offered product shall have at least one of the following certifications:
CE Certificate in accordance with 93/42/EEC Medical Device Directive.
US FDA Approval.
8. User Training
8.1 The supplier shall provide comprehensive user training covering:
Proper operation of the equipment.
Safety precautions.
Routine user maintenance procedures.
9. Warranty
9.1 The supplier shall provide a comprehensive warranty for a minimum period of two (2) years from the date of acceptance.
10. Maintenance Service During Warranty Period
10.1 During the warranty period, the supplier shall provide corrective and breakdown maintenance services whenever required.
11. Documentation
11.1 One (1) User (Operating) Manual in English.
11.2 One (1) Service (Technical/Maintenance) Manual in English.
11.3 A detailed list of important spare parts and accessories shall be provided, including:
Part numbers.
Unit prices/costing.
11 | Laryngoscope Set – Adult/Child
1. General Requirements
1. The set shall include both adult and pediatric laryngoscopes with appropriate blades for each age group.
2. The equipment shall be brand new, latest model (manufactured 2025 or later).
3. All components shall be autoclavable and suitable for repeated sterilization cycles without degradation.
4. The manufacturer shall have proven international experience in medical airway devices.
2. Regulatory & Standards Compliance
The equipment must comply with:
ISO 13485
ISO 10993 (biocompatibility)
EU MDR 2017/745 (CE Marking under MDR) or US FDA clearance
IEC 60601-1 (for battery-powered laryngoscopes)
All certificates and declarations must be submitted with the technical offer.
3. Laryngoscope Components
3.1 Handles
Material: Stainless steel or high-grade anodized aluminum
Ergonomic design for easy grip
Autoclavable
Compatible with both standard and pediatric blades
Power source: Rechargeable battery or disposable batteries (with at least 2 hours continuous illumination)
Switch: Reliable on/off toggle or push-button
Non-slip finish for safe handling
3.2 Blades
Adult Blades
Sizes: #3, #4 (Macintosh curved), #3, #4 (Miller straight)
Material: Stainless steel or high-strength autoclavable alloy
Light source: Fiber optic or LED illumination
Smooth, atraumatic edges
Pediatric Blades
Sizes: #0, #1, #2 (Macintosh and Miller options)
Material: Stainless steel or autoclavable alloy
Fiber optic or LED illumination
Smooth, atraumatic edges
3.3 Illumination
High-intensity LED light (minimum 50 lumens)
Consistent brightness throughout battery life
Backup light option in case of failure
3.4 Accessories
Rechargeable battery charger and spare batteries
Blade storage box or tray with separation for pediatric/adult blades
Protective blade covers or sheaths
Cleaning and sterilization instructions included
4. Operating & Storage Conditions
Operating temperature: 15°C – 35°C
Storage temperature: -5°C – 40°C
Relative humidity: 10–95% non-condensing
5. Documentation (Mandatory)
The supplier shall provide:
1. ISO 13485 certificate
2. CE MDR certificate or US FDA clearance
3. Declaration of Conformity
4. User (Operating) Manual – English
5. Service / Maintenance Manual – English
6. Cleaning, disinfection, and sterilization instructions
7. Spare parts list with part numbers
8. Warranty certificate
6. Warranty & After-Sales Service
Comprehensive warranty: Minimum 2 years
Preventive maintenance support
Spare parts availability: minimum 10 years
Service response time ≤ 72 hours
7. Acceptance Criteria
Final acceptance shall be based on:
Visual inspection for defects
Blade fit and handle compatibility
Illumination test
Battery performance check
Verification of submitted documentation
12 | Dental Unit
Description of Function
A Dental chair for Dental treatment
Operational Requirements
It shall operate on AC power supply.
Dental Unit with complete attachments and accessories.
Patient chair, 1 unit
Working stool, 1 unit
A built-in light cure unit
A built-in ultrasonic scaler with one each of pointed and flat scaler tips
One air compressor at least 1 horse power
Bidder shall indicate brand and model information here and provide technical data document for major components specified above.
Technical Specifications
Main unit standard configuration as follow:
Patient chair:
Electric chair movement and deluxe double articulated headrest.
With one left armrest as standard.
3 pre-set chair positions: start, treatment and rinsing position.
Electrical patient chair loading capacity: not less than 135kg.
Backrest movement range 105°-175°.
The lowest position of the patient chair from the ground shall not be less then 380mm.
The highest position of the patient chair from the ground shall not be less than 780mm.
Chair movement is controllable by the 4 way foot control at the chair base without touch panel.
No cables on the floor, hygienic and clean.
One main switch to control air, water and power.
The chair position is locked while an instrument is working.
With chair-backrest safety system, backrest and seat movement can be stopped once it meets obstacle.
Dentist element:
Dentist element with whip arm system.
Height of dentist element is adjustable.
1 X-ray film viewer (12V, 2000cd/m2).
1 silicon mat for the dentist element which can be sterilized.
1 three way syringe.
3 ISO 4-hole/Midwest hand piece hoses.
1 air pressure meter.
Assistant element:
1 three way syringe.
1 strong suction hose.
1 saliva ejector.
With suction filter system.
Water unit:
The cuspidor can be swivelled and removable for easy cleaning.
Cup filler and bowl rinsing systems shall prevent over filling of cup and prolong rinsing of bowl. Preferably programmable.
With automatic water heating system (24V)
With water venturi and air water separator system
Fresh water bottle, at least 1.5L
Operating light:
Colour temperature: 3800-4500K.
12V, 50 Watt Halogen bulb or Better.
Dental light intensity: min 25000lux with intensity dimming function.
Working stool, 1 units.
Mobile on 5 castors.
Height of seat and backrest is adjustable.
Backrest angle is adjustable and lockable.
Come with a built-in light cure unit.
Come with one suction unit
Come with a built-in ultrasonic scaler with one each of pointed and flat scaler tips.
Come with one air compressor at least 1 horse power or capacity sufficient to supply to the Dental chair specified above, whichever higher. Bidder shall indicate capacity of the unit offered here. Air compressor must be oil less & noise less
Accessories, Spare Parts and Consumables
All standard accessories, consumables and parts required to operate the equipment.
Power supply
Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 meters in length.
Standards & Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA approved product certificate.
Electrical safety conforms to standards for electrical safety IEC-60601.
User Training
The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
The warranty period for this item shall be two years
Installation and Commissioning
The bidder must be installed and commissioned equipment
Documentation
User (Operating) manual in English
Service (Technical / Maintenance) manual in English
Country Of Origin
Manufacture / Brand
Model / Type
Delivery time
Total Price (USD)
13 | Digital X-ray Machine
Description of Function
A general purpose digital X-ray machine for routine X-ray examinations at healthcare facilities.
Operational Requirements
It shall be suitable to be used for adult and pediatric patients in general Radiography examination.
System Configuration
X-ray Generator,1 unit
X-Ray tube & tube support system, 1 unit
Radiographic patient table, fixed height1 unit
Floor mounted bucky stand, 1 unit
Technical Specifications
Should be floor to floor mounted
X-ray Generator:
Microprocessor based, high frequency inverter generator, the generator shall have at least 100kHz.
Generator Output: not less than 30 Kw
Radiographic voltage: 40 kV to 140kV, in 1Kv step.
Radiographic current approx., 10 to 500 mA
mAs range: 0.5 mAs to 650 mAs.
Anatomical Programmable Radiographic mode shall be available.
Manual & automatic exposure control and automatic brightness control shall be available.
Shall come with overload protection device.
X-Ray Tube:
Maximum tube output shall match with the generator output
X-ray tube rotating: +/-90°.
Anode rotating speed: 2700 rpm or more.
Large focus not more than 1.2 mm.
Small focus not more than 0.6 mm.
Maximum tube voltage 140 KV.
Max housing heat dissipation rate: 250HU/s
Anode heat storage capacity: min 140 kHU (100kJ);
Max. heat storage capacity of the tube housing: 9,000kJ (1,250kHU).
Cooling method passive or forced air and/or oil cooling.
Filtration: min 2.7 mm Al equivalent.
Floor Mounted Tube Stand:
Horizontal travel range: 138 cm to 220 cm, movement arrested by electromagnetic brakes;
Vertical travel range: 150 cm to 160 cm movement arrested by electromagnetic brakes;
Central beam height: min 35 cm to min 185 cm;
Rotation of tube around vertical axis: ± 90°; stop positions: 0°, ± 90°;
X-ray tube rotation: ± 120°; stop positions: 0°, ± 90°.
Flat Panel Detector:
Dimensions (active field): min 34 cm x max 43 cm;
Detector thickness: 16 mm
Pixel size: 139 µm;
Spatial resolution (Nyquist frequency): 3.6 lp/mm.
Collimator:
Inherent filtration: min 1 mm Al at 70 kV;
Full-field light localizer: 150 W halogen light, timer;
Built-in light switch should be provided.
Copper prefilter: Without filter, 0.1 mm, 0.2 mm, 0.3 mm Cu;
Rotation: Up to a maximum of ± 45°;
Collimation control: Manual.
Floor Mounted Bucky Stand:
Vertical movement range: min 150 cm;
Vertical travel range: min 35 cm to min 185 cm;
Grid Pb 10/80; f0 = 115 cm and 150 cm/180 cm, stationary grid;
Front-cover-to-detector distance: < 40 mm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Automatic exposure control: Standard.
Cassette size: accept all sizes from 5”x7” to 14”x17”.
Movement arrested by electromagnetic brakes.
Patient Table:
Table height: min 72 cm;
Tabletop width: min 80 cm;
Tabletop length: min 220 cm;
X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
Tabletop travel: mon two panes, ± 46 cm, longitudinal, ± 10 cm, transverse;
Max. patient weight: min 200 kg;
Grid Pb 10/80; f0 = 115 cm, stationary grid;
Max. patient coverage: min 185 cm without patient repositioning;
Detector cover range (edge-to-edge): ≥ 55 cm, movement arrested by electromagnetic brakes;
Tabletop – detector distance: ≤ 70 mm;
Workstation (Standard configuration):
The imaging station is specially designed for diagnostic radiology and delivers excellent image quality with a fast and seamless workflow;
High-resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor, min 2 MP;
High performance PC included;
DR printer
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
Power supply:
Power supply: 415 ± 5%V (3 Phase), 50Hz
Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
The unit shall be capable of operating in ambient temperature of 20-40 deg C and relative humidity of 80%
Standards & Safety Requirements
Should be CE or FDA approved product certificate.
Must submit ISO13485
Training
Must provide user & service training.
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets (hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets (hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.