## Sheet1
1 | Donor Chair
No. | Item Specifications | Your Specifications
1 | Comfortable chair type with soft padding for cushioning and rexin cover.
2 | Seat, back rest and leg rest size designed for donor comfort.
3 | It should have height adjustment approx 58 – 60 cm.
4 | Adjustable arm rest for donor’s comfort and phlebotomist friendly
5 | Easily tilted to head low position, manual operated
6 | Comfortable working level for the operator.
7 | Load capacity approx 150 kg.
Operating Environment
1 | The system offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
Standards & Safety Requirements
1 | Must submit ISO 13485:2003/AC: 2007 AND
2 | CE or USFDA approved product certificate.
User Training
1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
1 | Comprehensive warranty for 2 years after acceptance.
Maintenance Service During Warranty Period
1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
Installation and Commissioning
1 | The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
Documentation
1 | User (Operating) manual in English.
2 | Service (Technical / Maintenance) manual in English.
3 | List of important spare parts and accessories with their part numbers and costing.
2 | Blood Collection Monitor
No. | Item Specifications | Your Specifications
1 | Should have a facility for gentle and uniform mixing of blood and anticoagulant.
4 | Should have facility to view the collection time
5 | Should have detachable tray for easy cleaning
6 | Should have motor activated clamping system and automatic clamping for low rate ,<20 ml/mt for more than 2 mts.
8 | Should have protection against Electrical shock.
9 | Oscillation details : 12+2 RPM, Motor driven
10 | Should have volume setting ranges from 50ml to 500m1 in step of 5 ml ,Automatic stronge and recall of set volume.
11 | Should have a LCD display with backlight.
12 | Accuracy : + 2% of programmed volume
13 | Should have the following alarm indications
14 | Indication and audible alarm for debit flow when flow rate goes below 20 ml/mt or high flow rate above 180 ml/mt.
15 | Indications and audible alarm at the end of collection .
16 | Indications & audible alarm during power failure, LED blinking when battery low.
17 | Indications and audible alarm during power failure.
18 | Power supply : 220/240 V AC , 50Hz
Operating Environment
1 | The system offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
Standards & Safety Requirements
1 | Must submit ISO 13485:2003/AC: 2007 AND
2 | CE or USFDA approved product certificate.
User Training
1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
1 | Comprehensive warranty for 2 years after acceptance.
Maintenance Service During Warranty Period
1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
Installation and Commissioning
1 | The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
Documentation
1 | User (Operating) manual in English.
2 | Service (Technical / Maintenance) manual in English.
3 | List of important spare parts and accessories with their part numbers and costing.
3 | Bench Tube Sealer
No. | Item Specifications | Your Specifications
1 | Should be a heavy duty tube-sealer capable of sealing tubes of various manufacturers of blood bag.
2 | Should be for bench-top .
3 | The sealing time should be adjustable between 0.5-5 seconds
4 | Sealing triggering should be automatic
5 | Should also have extended portable hand unit. Sealing hand should be with coaxial cable of 1.5 - 2.0 meter.
6 | Should have indication lamps for “Sealing Process” on handle as well as main unit.
7 | No warm-up time should be required.
8 | Should ensure easy separation of tube segments after the sealing.
9 | System should run on both mains and battery (more than 10 hrs. back up and charger).
10 | Should be lightweight.
11 | Detection of wet tube, leakage & sealing defect. Alarm in case of seal not complete.
12 | Power supply : 220/240 V AC , 50Hz
Operating Environment
1 | The system offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
Standards & Safety Requirements
1 | Must submit ISO 13485:2003/AC: 2007 AND
2 | CE or USFDA approved product certificate.
User Training
1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
1 | Comprehensive warranty for 2 years after acceptance.
Maintenance Service During Warranty Period
1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
Installation and Commissioning
1 | The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
Documentation
1 | User (Operating) manual in English.
2 | Service (Technical / Maintenance) manual in English.
3 | List of important spare parts and accessories with their part numbers and costing.
4 | Cold Centrifuge
No. | Item Specifications | Your Specifications
1 | Stainless steel housing
2 | Maximum speed around: 10,000 (RPM)
3 | Maximum RCF(FORCE): 20,000 RCF
4 | Capacity 12x450 500 ml,swing-out rotor
5 | Lid locking and holding
6 | Rotor: Number p-rotor 12
7 | Emergency lid lock release
8 | Motor overheating protection
9 | Imbalance switch-off
10 | Standstill indication
11 | Refrigerated centerifuge floor stand
12 | temp:controlable from -20c to +40c
13 | Input power supply: 3 phase 415 ± 20 % V AC , 50Hz
Operating Environment
1 | The system offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
Standards & Safety Requirements
1 | Must submit ISO 13485:2003/AC: 2007 AND
2 | CE or USFDA approved product certificate.
User Training
1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
Warranty
1 | Comprehensive warranty for 2 years after acceptance.
Maintenance Service During Warranty Period
1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
Installation and Commissioning
1 | The bidder must arrange for the equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
Documentation
1 | User (Operating) manual in English.
2 | Service (Technical / Maintenance) manual in English.
3 | List of important spare parts and accessories with their part numbers and costing.
5 | Manual Plasma Extractor
No. | Item Specifications | Your Specifications
1 | Description of Function
1.1 | Suitable for manually separating plasma from all types of blood collection bags.
2 | Operational Requirements
2.1 | Manual Plasma Extractor (Plasma Separation Stand) complete system.
3 | System Configuration
3.1 | Manual Plasma Extractor (Plasma Separation Stand) complete system.
4 | Technical Specifications
4.1 | Spring loaded bag holder. Ensure that bag stays in position.
4.2 | Compressor plate design to exert uniform pressure on donor bag minimizing chances of rapture of bag.
4.3 | Mode of operation – manual.
4.4 | Acrylic pressure / plate.
5 | Accessories, spares and consumables
5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
6 | Operating Environment
6.1 | The system offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
7 | Standards, Safety and Training
7.1 | Must submit ISO 13485:2003/AC: 2007 AND
7.2 | CE or USFDA approved product certificate.
8 | User Training
8.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
11 | Installation and Commissioning
11.1 | Supplier must accomplish proper installation & commissioning of the equipment on site.
12 | Documentation
12.1 | User (Operating) manual in English.
12.2 | Service (Technical / Maintenance) manual in English.
12.3 | List of important spare parts and accessories with their part numbers and costing.
6 | Blood Bank Refrigerator
No. | Item Specifications | Your Specifications
1 | Description of Function
1.1 | Blood Bank Refrigerator is used to store blood bags under controlled temperature.
2 | Operational Requirements
2.1 | System required with weekly chart recorder and digital display.
3 | System Configuration
3.1 | Blood Bank Refrigerator with weekly chart recorder, digital display and with complete accessories.
4 | Technical Specifications
4.1 | Temperature range:
Must have adjustable temperature control range from +2 to +6 . It shall maintain internal temp at 4 oC & the uniformity of this temperature must be maintained throughout the chamber with the maximum average temperature variation of +/- 1 oC between different chambers.
4.2 | Capacity must not less than 120 litres. (not less than 60 Blood bags 450ml).
4.3 | Blood Bank Refrigerator shall have integrated temperature monitoring system with microprocessor controls.
4.5 | The blood bank refrigerator dpuble door shall have a large LCD which displays:
Temperature
High & Low alarm points with date and time
Previous 24 hours temperature in graphical from Date of power failure/resumption in last 24 hours with date & time
4.6 | The blood bank refrigerator shall also have an inbuilt circular chart recorder for 7 days recording of temperature on circular chart paper.
4.7 | The internal automatic temperature alarm system shall work if a temperature falls below 2 oC & exceeds beyond 6 oC.
4.8 | The internal temperature alarm system shall also have a battery backup of minimum 3-4 hours.
4.9 | Internal construction must be made up of high grade stainless steel 304 (min 22 G). External construction Corrosion resistant sheet at least 1 mm thickness.
4.1 | It shall have lockable door. Outer door shall be made of glass to see through and inner door shall be made of acrylic sheet to ensure ease of operations, better maintenance of internal temperature.
4.11 | Blood Bank Refrigerator shall confirm to noise level of less than 85 dBA as per IEC 61010.
4.12 | Internal cabinet lighting to be provided with lamp illumination whenever door opens.
4.13 | Shall come with roll out steel trays for proper storage of blood bags.
4.14 | Blood bank refrigerator shall have in built servo controlled voltage stabilizer of suitable rating.
4.15 | Shall have adjustments for uneven bases. The adjustments must be easy to use like rotating a screw at the legs in the base.
5 | Accessories, spares and consumables
5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
6.2 | Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
6.3 | Suitable Automatic Voltage regulator/stabilizer meeting international standards must be supplied. Broad specifications are: Automatic Type Input 150-280V, Output 220 V +/- 7 %, 50 Hz. Single phase, AC with automatic 2-4 sec Cut Off and 6-9 minutes restart delay. Quick start arrangements for bypassing the start delay. Suitable MCB on input voltmeter and indicators on Front Panel. Input Pore Cable with 15 A Plug and six way output terminal strip for two outlets
7 | Standards and Safety Requirements
7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 AND
7.2 | Must comply with WHO/UNICEF Specification Reference: BTS/RF.1 and DIN 58371.
7.3 | Test and inspections as per WHO Procedure reference: Laboratory Test Procedure: Standard Test Procedure: BTS/Proc. / 3.
7.4 | Shall meet IEC 60335-1and -2-24 General requirements of electrical safety.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure planned preventive maintenance (PPM) and corrective/breakdown maintenance whenever required.
11 | Installation and Commissioning
11.1 | The supplier must accomplish proper commissioning of the item onsite.
12 | Documentation
12.1 | User (Operating) manual in English
12.2 | Service (Technical / Maintenance) manual in English
12.3 | List of important spare parts and accessories with their part number and costing.
7 | Ultra Plasma freezer (-40 °C)
No. | Item Specifications | Your Specifications
1 | Description of Function
1.1 | Deep Freezers are required to preserve blood and blood products, vaccinations etc at specified temperature.
2 | Operational Requirements
2.1 | Microprocessor Vertical Freezer, at least double door with adjustable 4 to 6 shelves ( frost free)
2.2 | Separate Chamber racks to be pulled out for easy handling
2.3 | Non-CFC refrigerant
3 | Technical Specifications
3.1 | Capacity within 250 to 300 L
3.2 | Digital display of set and actual temperature, with audiovisual alarm
3.3 | No condensation on storing material with automatic electric defrost
3.4 | Construction:
Solid rust free cabinet to prevent corrosion and lockable castor wheels.
3.5 | Refrigeration System
Heavy Duty refrigeration system, maintenance free, below -40 deg C (+ 10 C) cascaded connection with hermetically sealed refrigeration compressors and reliable refrigeration to minimize noise and vibration, air cooled with security lock to prevent unintentional switch off shall be supplied. It should have cooling time within hours at maximum ambient temperature of 33deg C.
The equipment should be of continuous duty and frost free.
3.6 | Alarm
It should also have audio visual Electronic Alarm System independent of power supply.
3.7 | Insulation High density polyurethane or equivalent Gaskets - Double seal silicon.
4 | System Configuration Accessories, spares and consumables
4.1 | As specified
5 | Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -50deg C and relative humidity of 15-90%
The unit shall be capable of operating continuously in ambient temperature of 10 -40deg C and relative humidity of 15-90%
6 | Power Supply
6.1 | Power input to be 220-240VAC, 50Hz fitted with Sudanese plug
6.2 | Suitable Servo controlled Stabilizer/CVT
6.3 | Resettable over current breaker shall be fitted for protection
7 | Standards and Safety
7.1 | Should be FDA or CE or ISI approved product
7.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
7.3 | Should be compliant to ISO 13485: Quality systems - Medical devices - Particular requirements for the application of ISO 9001 applicable to manufacturers and service providers that perform their own design activities.
8 | Documentation
8.1 | User manual in English
8.2 | Service manual in English
8.3 | List of Equipments available for providing calibration and routine Preventive Maintenance Support. As per manufacturer documentation in service/technical manual.
8.4 | Log book with instructions for daily, weekly, monthly and quarterly maintenance checklist.
8.5 | List of important spare parts and accessories with their part number and costing
8 | Ultra Plasma freezer (-80 °C)
No. | Item Specifications | Your Specifications
1 | Description of Function
Deep Freezers are required to preserve blood and blood products, vaccinations etc at specified temperature.
2 | Operational Requirements
2.1 | Internal Minimum Capacity 350 to 400 L net at least double door with adjustable At least 6 shelves
2.2 | Range up to -65C to -85C(Adjustable)
2.3 | Vertical Cabinet (upright model)
3 | Technical Specifications
3.1 | Construction:
Solid rust free cabinet to prevent corrosion and lockable castor wheels. Inner surface should be stainless steel.
3.2 | CONTROL SYSTEM
Micro-processor based temperature controller with digital temperature display LED/LCD with seven days graphic
temperature recorder with rechargeable battery back up including charger maintenance free and insensitive to vibration. Details of battery and battery charger shall be indicated.
3.3 | Refrigeration System
Heavy Duty refrigeration system, maintenance free, below -80 C (+ 10 C) with hermetically sealed dual compressors, noise free and vibration free, air cooled with security lock to prevent unintentional switch off shall be supplied. It should have maximum cooling time of 5 hours at maximum ambient temperature of 330C. The equipment should be of continuous duty.
3.4 | Alarm
It should also have audio visual Electronic Alarm System independent of power supply.
3.5 | Insulation
High density polyurethane or equivalent Gaskets - Double seal silicon.
3.6 | Door heating system for easy opening of door
4 | System Configuration Accessories, spares and consumables
4.1 | Refrigerator -01
5 | Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -50deg C and relative humidity of 15-90%
The unit shall be capable of operating continuously in ambient temperature of 10 -40deg C and relative humidity of 15-90%
6 | Power Supply
6.1 | Power input to be 220-240VAC, , 50Hz,/440 V 3 Phase as appropriate fitted with Sudanese plug
6.2 | Suitable Servo controlled Stabilizer/CVT
6.3 | Resettable over current breaker shall be fitted for protection
7 | Standards and Safety
7.1 | Should be FDA or CE or ISI approved product
7.2 | Electrical safety conforms to standards for electrical safety IEC-60601 / IS-13450
7.3 | Should be compliant to ISO 13485: Quality systems - Medical devices - Particular requirements for the application of ISO 9001 applicable to manufacturers and service providers that perform their own design activities.
8 | Documentation
8.1 | User manual in English
8.2 | Service manual in English
8.3 | List of Equipments available for providing calibration and routine Preventive Maintenance Support. As per manufacturer documentation in service/technical manual.
8.4 | Log book with instructions for daily, weekly, monthly and quarterly maintenance checklist.
9 | ELISA Reader and Washer
No. | Item Specifications | Fill your Specification
1 | Description of Function
1.1 | ELISA Reader is required to Read the Color Density known as OD (Optical Density) in ELISA (Enzyme Linked Immuno-Sorbent Assay.) Plates.
1.2 | A washer for microtitre plates designed to ensure thorough washing of reagents between Enzyme-Linked Immunosorbent Assay (ELISA) steps.
2 | Operational Requirements
2.1 | ELISA Reader and washer 8 channles complete with Printer is required.(Plate reader 96).
3 | Technical Specifications for Reader
3.1 | Should have 8-12 measuring channel & reference channel
3.2 | Should have wave length range of 340- 750 nm 6 filters 340, 405. 450, 492, 540, 630nm with provision for fitting any additional filters
3.3 | Should have an absorption range of 0-4.000A
3.4 | Should have a resolution of 0.001A
3.5 | Should read within 6-8 seconds
3.6 | The control panel should have soft color touch screen display, capable of showing graph etc.
3.7 | Should have external & internal programmable time & speed shaking
3.8 | Should be able to read all types of plates
3.9 | Should have a single halogen lamp with save features as light source
3.1 | Should have user defined programs 30 or more.
3.11 | RS232/USB output for Printer, PC connectivity and Data acquisition should be there
3.12 | Should have data memory of 300 plates.
3.13 | Should have external printer, capable of printing complete results & graphs etc. from Elisa system
4 | Technical Specifications for Washer
4.1 | 8-channel strip manifold, open system.
4.2 | Rinse & prime programme.
4.3 | Wash parameters include: 16-character assay name, number of cycles, wash volume, flow rate and variable soak times.
4.4 | Dispense only and aspirate only modes for reagent addition and removal.
4.5 | Shall have built-in multi-speed shaker for improved CVs and reduced assay backgrounds.
4.6 | Shall have crosswise aspiration/double aspiration of flat bottom micro-plates for reduced residual liquid.
4.7 | Bottom wash mode for rapid dilution of reagent.
4.8 | Shall have built-in vacuum & pressure pump assembly.
4.9 | Bottles for waste rinse and wash.
4.1 | Accommodates flat, U or V-shaped bottom plates.
4.11 | Wash cycles: Between1-10.
4.12 | Dispensing volumes from 25 to 3000ul.
4.13 | Soak time 1-600 seconds.
4.14 | Fluid flow rate in 150 to 1000ul/well/second to accommodate cellular assays.
4.15 | Spill-over protection & electronics isolated from fluidics.
4.16 | Optional automatic buffer switch in flip out aerosol cover or similar.
5 | System Configuration Accessories, spares and consumables
5.1 | System as specified.
5.2 | Halogen Lamps : 2
5.3 | External Printer
5.4 | Dust Cover -01
5.5 | Set of pipettes consisting of single channel variable volume color pipettes 0.5-10 ul, 5-40 ul, 40-200 ul,200-1000 ul
5.6 | 8 channel variable volume color multi-channel pipettes 5-50 ul and 50-300 ul.
6 | Accessories, spares and consumables
6.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
7 | Operating Environment
7.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
7.2 | Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
7.3 | Suitable UPS with maintenance free batteries, voltage regulation and spike protection for minimum 30 min. back-up shall be supplied with the system.
8 | Standards and Safety Requirements
8.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
8.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
8.3 | Must comply with IEC 61010 Safety requirements for electrical equipment for measurement, control, and laboratory use
9 | User Training
9.1 | Must provide user training (including how to use and maintain the equipment).
10 | Warranty
10.1 | Comprehensive warranty for 2 years after acceptance.
11 | Maintenance Service During Warranty Period
11.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
12 | Installation and Commissioning
12.1 | Supplier must accomplish proper installation and commissioning of the equipment on site.
13 | Documentation
13.1 | User (Operating) manual in English
13.2 | Service (Technical / Maintenance) manual in English
13.3 | List of important spare parts and accessories with their part numbers and costing.
10 | Platelet Shaker with Incubator
No. | Item Specifications | Fill your Specification
1 | Technical Specifications
1.1 | Stainless steel chamber with adjustable shelves and a tough ended glass inner viewing door.
1.2 | The outer cabinet is to be rust resistant.
1.3 | Temperature Control detail required:-
1.4 | An LED display to show the chamber temperature, Indicator
1.5 | Lamps to show when the heater is active and if an over temperature condition exists.
1.6 | The over temperature safety cut-out to be set by the user.
1.7 | Fitted with circulation fan.
1.8 | Temperature Range : At least 5°C above ambient to +60°C
1.9 | Control (fan) : ±0.1°C at +37°C
1.1 | Variation (fan) : ±0.25°C at +37°C
1.11 | Chamber Capacity: 80 - 100 Litres
1.12 | Shelves: >= 5.
2 | Operating Environment
2.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
2.2 | Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
2.3 | Suitable UPS with maintenance free batteries for minimum one-hour back-up should be supplied with the system.
3 | Standards and Safety Requirements
3.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
3.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
4 | User Training
4.1 | Must provide user training (including how to use and maintain the equipment).
5 | Warranty
5.1 | Comprehensive warranty for 2 years after acceptance.
6 | Maintenance Service During Warranty Period
6.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
7 | Installation and Commissioning
7.1 | Supplier must accomplish proper installation and commissioning of the equipment on site.
8 | Documentation
8.1 | User (Operating) manual in English
8.2 | Service (Technical / Maintenance) manual in English
8.3 | List of important spare parts and accessories with their part numbers and costing.