## 01-2025
Electroconvulsive Therapy Machine (ECT)
No. | Technical Specifications | Fill Your Specifications
1 | Description of Function:
The main aim of Electroconvulsive Therapy is to cause a massive convulsion in the brain (a massive epileptic fit). This is achieved by giving the brain an electric shock using an ECT Machine. ECT machines are, basically, transformers which modify Mains Current so that it is transmitted to the patient's skull in timed pulses.
2 | Operational Requirements
1. The unit should have Parameter Display on LCD/LED
2. Should have auto Stimulus Voltage
3. Should have Auto Impedance Check
4. Should provide Output Display in joules & milli coulombs.
3 | Technical Specifications
Technical data stimulus:
1. Bidirectional Square Wave
2. Current: 0.8 Amp Constant
3. Frequency Range: 70 Hz (Fixed)
4. Pulse Width: 1 ms (fixed)
5. Mode: Auto & Manual
6. Stimulus Duration in Auto mode: 0.1 to 5.9 sec. in step of 0.1 sec.
7. Cerebral Stimulation 0 to 40 volts
4 | 5. Power Supply
1. Power input to be 220-240VAC, 50Hz
2.UPS of suitable rating with voltage regulation and spike protection for 60 minutes back up
5 | Accessories, spares and consumables:
1. All the accessories to make equipment functional as per specifications should be supplied.
2. All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment:
1. The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
7 | Standards and Safety Requirements:
1. Must submit ISO13485:2003/AC:2007 for Medical Devices AND
2. CE (93/42 EEC Directives) or USFDA approved product certificate.
8 | User Training:
1. Must provide user training (including how to use and maintain the equipment).
9 | Warranty:
1. Comprehensive warranty for 2 years and must replace all the parts as and when required.
2. Maintenance Service During Warranty Period
3. During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
10 | Documentation:
1. User (Operating) manual in English.
2. Service (Technical / Maintenance) manual in English.
3. List of important spare parts and accessories with their part numbers and costing.