## Sheet1
RFQ 04-2025- Specs
1 |  | Full Automation Coagulation Analyzer with full accessories and with UPS
| No. | Item Specifications | Fill your Specification
| 1 | Technical Specification
Bench Top
Open system ( at least for PT and APTT)
Able to read Clotting , Chromogenic and Immunological Assay
Volume of sample not more than 50 ul and reagent not more than 100
QC program
Easy work
Capable of doing all factor assay,coagulation test including thrombophilia screening
Random Access
Liquid level sensor
Able to do more than 4 different tests per run
Run should be not less than 10 samples
At least 2 probes
Must be supplied with computer and all accessories
| 2 | Operating Environment
The unit shall be capable of operating continuously in ambient temperature of 10 -40deg C and relative humidity of 15-90%.
The unit shall be capable of being stored continuously in ambient temperature of 0 -50deg C and relative humidity of 15-90%.
| 3 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
Suitable UPS with maintenance free batteries for minimum 30 min. back-up shall be supplied with the system.
| 4 | Standards and Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
Should be FDA/CE/BIS approved product.
Shall meet internationally recognised for Electromagnetic Compatibility(EMC) for electromedical equipment: 61326-1.
Certified to be compliant with IEC 61010-1, IEC 61010-2-281, 61010-2- 101 for safety.
| 5 | User Training
Must provide user training (including how to use and maintain the equipment).
| 6 | Warranty
Comprehensive warranty for 2 years from acceptance.
| 7 | Maintenance Service During Warranty Period
During the warranty period supplier must ensure planned preventive maintenance (PPM) and corrective/breakdown maintenance whenever required.
| 8 | Installation and Commissioning
The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the user in advance, in detail.
| 9 | Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts and accessories with their part number and costing.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
2 |  | Blood Sample Rolling Mixer
| No. | Item Specifications | Fill your Specification
| 1 | Technical Specification
Rocking and rolling action for complete mixing
Removable rollers to accommodate large bottles
Suitable for continuous operation
Nomber of Rollersapprox: 6 .
Speedrpm approx: 33 .
Amplitudeapprox: 16 mm .
Roller Approx Size: 340x30 mm
Input power supply: 220/240 volt AC , 50Hz Schuko plug
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
3 |  | Binocular Microscope (Electric)
| No. | Item Specifications | Fill your Specification
| 1 | Description of Function
| 1.1 | Compoun d microscope consists of two or more than two magnifying lenses. One can view individual cells, even living ones. It has high magnification
| 2 | Operational Requirements
| 2.1 | System co mplete with illumination system is required.
| 3 | System Configuration
| 3.1 | Binocular Microscope Compound with complete accessories
| 4 | Technical Specifications
| 4.1 | Body :Bin ocular, sturdy, stable base body with focus adjustment controls
| 4.2 | Eye piece : Paired, high quality, (the image of the object as seen through the binocular eyepiece must be well defined centrally in at least 2/3 field of view), achromatic, wide field, 10x and 15x without inbuilt pointer. The eyepiece must be aplanatic and have a minimum field number of 18. Dioptre adjustment must be present on one/ both eye pieces or on the eye piece tube
| 4.3 | Objective : Four 4x, 10x, 40x, 100x.
| 4.4 | 10x and 4 0x objectives must have numerical apertures of 0.25 and 0.65 respectively and must be of spring loaded type or otherwise.
| 4.5 | 100 x mu st have numerical aperture of 1.25 and must be of oil immersion and spring loaded type.
Suitable prominent marking must be provided on 100x for easy identification.
| 4.6 | Unbreaka ble containers to be provided for storing the objectives. All objectives must be wide field, achromatic and par focal.
| 4.7 | Making f or the Objectives : Each objective must be engraved with the following information:-
| 4.8 | Nose piece: Revolving nose piece to accommodate four objectives with click stops. It must be provided with ribbed grip for easy rotation mounted on a precision ball bearing mechanism for smooth and accurate alignment. Extra ports if any must be fitted with dust proof metallic/ebonite caps.
| 4.9 | Stage Un iformly horizontal, mechanical stage having dimensions of length 140 mm (+/- 20mm) with fine vermier graduations (minimum reading accuracy of 0.1 mm). the stage must be provided with spring loaded slide holder for exact positioning of specimen/ slide. It must be designed with convenient sub-stage vertical coaxial adjustment for slide manipulation. The stage must have ball- bearing arrangement to allow smooth travel in transverse directions i.e. 80 mm (+/-5mm) and front to back direction, 50mm (+/-5mm)
| 4.1 | Sub-stage condenser: Abbe-type condenser, numerical aperture (N.A.) 1.25 focusable with rack and pinion arrangement incorporating a spherical lens and an iris-diaphragm. The condenser must have a filter holder and removable/ swing in/ out blue filter (suitable for bright field Microscopy).
| 4.11 | Sub-stage illuminator: 1.The system must have a build-in variable light source (Illuminator). This source must have a 20 W, 6/12 V Halogen lamp. The circuitry for the light source must include a constant voltage supply. The system must be provided with a step down transformer and an on-off switch and intensity control. The lamp must be provided with a lamp socket which has the facility for easy replacement of the bulb. light
| 4.12 | The Illuminator must have a build-in field diaphragm for Kohler illumination.
| 4.13 | Eye piece tubes: Binocular eye piece tubes, inclined at 30 and 45 degrees, rotatable through an angle of 360 degrees, having inter-pupillary distance range of 54-74 mm or wider, covering the above mentioned range
| 4.14 | Focusing knob: Co-axial coarse and fine focusing knobs capable of smooth fine focusing movement over the full range of coarse travel. The fine focusing movement must have sensitivity of two microns or less (finer) over the entire coarse focusing stop safety arrangement must be provided.
| 4.15 | General 1 .All optical parts including objectives, eye pieces and prisms must have anti-reflective coating which also gives anti-fungal property.
(100 ml) must be provided with each microscope.
purpose.
and not fixed on the Microscopes.
| 5 | Accessories, spares and consumables
| 5.1 | Accessories:
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plug to meet purchaser’s country requirements. The power cable must be minimum 3 metres long.
| 6.3 | Voltage corrector/stabilizer of appropriate ratings meeting international standards.( Input 160-260 V and output 220-240 V and 50 Hz)
| 7 | Standards and Safety Requirements
| 7.1 | Must sub mit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE or US FDA approved product certificate.
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years after acceptance.
| 10 | Maintenance Service during Warranty Period
| 10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 11 | Installation and Commissioning
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
| 12 | Documentation
| 12.1 | User (Operating) manual in English
| 12.2 | Service (Technical / Maintenance) manual in English
| 12.3 | List of important spare parts and accessories with their part numbers and costing.
| 12.4 | Certificate of calibration and inspection from factory.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
4 |  | Vortex Mixer
| No. | Item Specifications | Fill your Specification
| 1 | Technical Specification
A suspension mixer enabling controlled and reproducible conditions to be maintained.
Suitable for Coombes, serological, agglutination tests and mixing blood.
Speed 10-100 rpm
Angle to tilt 0-15 deg continuously variable
Power supply: 220/240 volt VAC , 50Hz Schuko plug. The power cable must be at least 3 metres in length.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
5 |  | Echo Ultrasound Machine with probes for Adult and Pediatric and with printer and UPS
| No. | Item Specifications | Fill your Specification
| 1 | Description of Function
| 1.1 | Colour Doppler Echocardiography System is required to study the anatomical abnormalities and blood flow in the heart and associated vessels.
| 2 | Operational Requirements
| 2.1 |  Latest generation Electronic Phased array Colour Doppler system with minimum 512 electronic independent channels.
 System must be DICOM ready and capable of being interfaced with HIS/RIS/PACS.
 Must be upgradable to next generation system on site.
 Frequency compounding or better technology for better resolution and penetration.
| 3 | System Configuration
| 3.1 | Colour Doppler System with all application packages, quad loop for serial studies with high frame rate review. Harmonic imaging capability in all modes. (Tissue, Contrast, Anglo) Integrated Stress Echo Package, Digital Storage and Retrieval – 01no.
 1-3 MHz Adult Cardiac probe Electronics Phased Array probe- 01 ea.
 3-11 MHz Electronics Phased Array Probe for Vascular applications- 01 ea.
 Multi-plane TEE Probe: 4-8 MHz for Adult as well as Paediatric echocardiography.
 5-10 MHz Electronic phased array probe for Paediatric cardiology.
 Colour Printer -01no.
 B/W Video Thermal Printer -01no.
| 4 | Technical Specifications
| 4.1 | Latest generation Electronic Phased array colour Doppler system with Minimum 512 Electronic independent channels.
| 4.2 | 256 gray shades for sharp contrast resolutions
| 4.3 | Multi-dimensional Beam former for generating two images simultaneously-one at low end of bandwidth and one at high end-then selectively retrieves and mixes the components together for finely textured 2-D or B mode image with superior contrast resolution
| 4.4 | Adult Cardiac and Vascular Probes to be supplied which must be latest generation wide band transducers without frequency selection for higher sensitivity of response over a broader frequency range of operation. All probes to be phased array.
Probes for paediatric application and Trans oesophageal Echo for future requirement.
| 4.5 | Harmonic Imaging: System must have following modes in harmonic with separate setting for:
 Tissue Harmonic.
 Contrast Harmonic - both triggered and real time
 Harmonic Angio.
 Quantification of harmonics imaging
| 4.6 | Harmonic imaging capability in Adult Cardiac, Paediatric Cardiac and linear Probe.
| 4.7 | Gain control in two dimensions for additional level of flexibility to image quality control.
| 4.8 | Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes
| 4.9 | Frame rate must be 300 FPS or more.
| 4.1 | Steerable PW/CW in all Phased Array probes.
| 4.11 | High definition acoustic zoom for enlarging sections of 2D and colour flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate.
| 4.12 | Modes - 4D, M-Mode, Steerable PW/CW Doppler, Colour Doppler, and High Definition colour flow with capability of automatically picking up colour flow as a function of focal depth
| 4.13 | Monitor must be 15" or more, high resolution colour monitor.
Tilt and Swivel monitor must be able to view in all angles and all light conditions.
| 4.14 | Colour Flow Imaging for:
 Increased lateral & spatial resolution.
 Detection of even subtle areas of turbulence, displaying a more physiological blood flow appearance without loss of frame rate.
 Colour flow with capability of automatically picking up colour flow as a function of focal depth
| 4.15 | Tissue Colorization (B-colour) for improved contrast resolution
| 4.16 | Application software for Adult, Paediatric, Foetal and Peripheral Vascular and Trans oesophageal applications. (All application packages must be built into the system).
| 4.17 | Cine loop memory- more than 120MB of memory.
 High Frame rate review for better clarity of playback images study in slow motion.
 Quad loop with memory for pre and post image comparison of any procedure.
 Memory: 256 frames or more in quad loop. M Mode & Doppler Scroll Memory-40 seconds or more.
 Frame grabber facility for post analysis.
| 4.18 | Various maps for pre and post processing.
| 4.19 | ECG triggers facility.
| 4.2 | User defined system and application pre-sets for multi-user department.
| 4.21 | Minimum 4.8 GB optical disc drive for image storage and retrieval. (standard with system)
| 4.22 | Dedicated integrated dynamic stress echo package for flexible user defined protocols with stacked sub loops facility and contrast stress protocol.
| 4.23 | Tissue movement colorization with quantification possibility for IHD/CAD patients.
| 4.24 | Three transducer ports will be preferred.
| 4.25 | Colour Map resolution up to 128 levels.
| 4.26 | Study Manager (> 1.5 GB) for on-cart digital acquisition, review and editing of complete patient studies.
| 4.27 | Facility of Real time perfusion studies
| 4.28 | System Peripherals shall include:
 CD Writer with calculation facility on playback.
 Colour Video Printer.
 B/W Thermal Printer.
| 5 | Accessories, spares and consumables
| 5.1 | Accessories:
 DVD/CD Recorder with 100 CDs and 100 DVDs
 Colour Print Paper- 500 sheets
 B/W Thermal Paper - 10 rolls
 ECG Cable - 02nos.
 MO Disc - 10pcs
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
| 6.3 | UPS of suitable rating conforming to international standards shall be supplied for minimum 30 min. backup for the entire system.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 7.3 | The product shall comply to IEC 60601-2-37 ed1: Medical Electrical Equipment - Part 2-37: Particular Requirements for the Safety of Ultrasonic Medical Diagnostic and Monitoring Equipment.
| 7.4 | Type of protection against electric shocks - Class I
Degree of protection against electric shocks for ultrasound probes Type "BF"
For ECG electrodes Type 'CF"
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years.
| 10 | Maintenance Service during Warranty Period
| 10.1 | During the warranty period supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 11 | Documentation
| 11.1 | User (Operating) manual in English
| 11.2 | Service (Technical / Maintenance) manual in English
| 11.3 | List of important spare parts and accessories with their part numbers and costing.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
6 |  | Stress ECG Machine Complete with Treadmill and with software with UPS
| No. | Item Specifications | Fill your Specification
1. Description of functions:
In this system, the patient exercises on a treadmill according to a standardized protocol and the cardiac abnormalities can be studied under stress conditions which we may miss under resting
2. Operational requirements
2.1 The treadmill stress test system should be complete with acquisition of resting and stress ECG, Treadmill unit with interface with all the protocols and provision of printing the resting as well as stress ECG and analyzing the same
2.2 Should be able to be interfaced to hospital information systems / LAN/WLAN
3. Technical specifications
1. System should have telemetry module (Blue tooth based) between patient and machine
2. USB connectivity for high speed data transfer from acquisition device to PC.
3. Should have integrated automatic programmable stress test NIBP measurement module with Bi-directional data transfer (compatible with the stress test system)
4. Repairable and generic patient cables and TMT/ECG accessories should be used
5. Should have software inbuilt online support system for controlling system through internet for providing online support.
6. Company should make spares available for the entire life of the system.
7. Service response time should be quick and service personnel should be able to attend the call within 48 hours (working days).
8. Quality certifications like ISO-9001:2008, ISO-13485:2012 or equivalent
9. Should have certification from USFDA and/or valid CE certificate from notified body.
Specification for Acquisition Software
1. Should acquire data from 12 Lead simultaneously
2. Should have frequency response from 0.05 – 150 Hz
3. Should have notch filter around 50 Hz
4. Should have sampling rate of at least 4500 samples/Second
5. Should have rejection ratio (CMRR) more than 100 dB
6. Should have internal Defib protection
7. Should have less than 10 uAmp patient leakage
8. Should have input impedance of more than 100 MOhms
Specification for stress test software
1. Should be capable of displaying real-time or stored ECG tracings.
2. Should display and regularly update ECG 12 leads, 12 medians, 1 expanded median, HR, BP, METS, Stage time, test time, protocol name, stage name, speed and grade of treadmill.
3. Should be able to give standard ECG report including standard ECG analysis.
4. Should have automatic stage print out facility at the end of each exercise
5. Median complexes of all 12 leads with rhythm leads with different formats should be available
6. Should have capability to display real-time ST running trend
7. Should have ability to display trend graph for HR, BP, ST level, ST slope and J amplitude
8. Should have automatic detection, display, storage and review of rhythm events
9. Should be able to display ECG in various formats like 3 Lead + 12 Median; 6 Lead + 12 Median; 12 Lead + 12 Median
10. Should have controllable display sweep speeds ranging from 5.0 to 100 mm/Sec
11. Should have display sensitivities ranging from 0.25 to 8.0 cm/mV
12. Should have base line correction (BLC) for stable baseline during test
13. Should run various test protocols like Bruce, Modified Bruce, Balke, Ellested, Naughton and user defined protocols.
14. Acquisition and analysis softwares should be upgradable to latest version free of cost.
15. Should have capability to import patient data from HIS and also manually edit/add data of patient
16. Raw data from software should be made available in standard formats for further analysis with softwares like MATLAB.
17. Should be able to print report in PDF format.
18. Should able to print reports with standard Laser printers on A4 Plain sheets.
Specifications for treadmill
Treadmill should have:
1. Fully interfaced and controllable from software
2. Controllable speed of 9-20 Km/H
3. Variable inclination (grade) from 0 – 22%
4. Adequate walking area ~ 1500 mm X 500 mm
5. Controlled by optically isolated RS 232 or USB
6. AC motor capable of running on inverter/UPS
7. Emergency stop feature
8. Power requirement 230 V,50 Hz, 15 A
Specification for computer
System should be supplied with Trolly mounted branded all in one PC which should have minimum:
1. i5 processor, 16 GB RAM, 2TB Hard Disk, DVD/CD Drive, ~ 23” touch screen LED display.
2. Windows 11 (64 bit) Operating System.
3. Laser Printer for printing on A4 Sheets.
4. 6 KVA online UPS for entire system backup for at least one hour.
Should be supplied with required standard accessories:
1. Disposable standard chest electrodes - 100 Pcs
2. Reusable ECG electrodes – 2 sets
3. Extra Patient cable - 01
4. Spike protector (with 5 sockets) - 01 No
5. Earthing cord – 1 No
6. L-N Key set
7. ECG jelly – 2 No
8. A4 Sheet Ream – 1 No.
9. Fuse (1.5 Amp and 10 Amp) – 10 each or compatible
10. User Manual – 01 No.
11. Cuff motion sensor – 01 No.
12. NIBP Module – 01 No.
13. Wrist strap of NIBP cable – 01 No.
14.Must submit ISO13485:2003/AC:2007 for Medical Devices ANDCE or US FDA approved product certificate.
15.Comprehensive warranty for 2 years after acceptance.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin: