## RFQ 5-2025
RFQ 05-2025- Specs
1 |  | Electric ICU Bed Five Movement With Mattress
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
Technical Specifications
1. Bed should have Dual Pedestal Design.
2. Should have built in Bed extender of 200mn for long patients.
3. Tuck away side rails with embedded controls.
4. Electrically operated controls should have lock out facility for electric controls at foot end
5. Backrest & leg section emergency release (CPR)
6. Should have Trendelenbug & covered with bed cover made of AB for the frequent disinfections and washing
7. Synchronized command for bed sprung bas with inclination - backrest – 75 degree and above & Femur rest 25-30 degree
8. Should have electric backrest and pelvic sections with simultaneous movement
9. Should have complying to international standards with acimulated power 1,5 degree 70 VA, electrical current protection class 1, shock protection level type B.
10. Release lever for the side rails with four sectors
11. Protection against water invasion
12. Should have following parameters:
a. Length: > 210 cm
b. Width: > 85 cm
c. Height adjustment: 40-85 cm
d. Back section 0-90
e. Leg Section: 0-25
f. Trendlenberg: 10-20
g. Reverse Trendlenberg: 5-10
13. Should have collapsible /tuck – away side rails
14. Should have double X-ray back rest with Cassette tray
15. Should have facility to fix IV rod at all the four corners and place for fixing accessories
16. Should have castors with central braking system and steering facility
17. Should be radiolucent from Neck though Pelvic region
18. Should have bumpers at all 4 corners
19. Should be provided with:
a. Bed ends fixed to the frame
b. Side rails
c. IV rods
d. Mattress
e. X-ray cassette tray, urine bottle holder, drainage bottle holder
20. Load bearing capacity for the bed should be at least 220kg or more, should have 90 degree upright backrest for chest Imaging.
21. Electrical safety conforms to standards for electrical safety IEC-60601/IS-13450
22. Should have Dual Castors of 18-20 cm
23. Electric shock protection facility
24. Zone controls in Bed to track movement of patient on Bed (Optional)
Operating Environment
The unit shall be capable of operating in ambient temperature of 15°C-35°C and relative humidity of 10-95%
The unit shall be capable of being stored continuously in ambient temperature of -5°C -40°C and relative humidity of 10-95%
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metres in length.
Standards & Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
Training
Must provide user & service training.
Warranty
Comprehensive warranty for 2 years.
Maintenance Service During Warranty Period
Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
2 |  | Diathermy Machine with Trolley
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of function
| 1.1 | Diathermy Machines are used for surgical cutting and for controlling bleeding by causing coagulation (haemostasis) at the surgical site
| 2 | Operational Requirements
| 2.1 | It shall operate on AC power supply in the operating theatre.
| 3 | System Configurations
| 3.1 | Diathermy Machine (Electrosurgical unit) 300W with complete accessories.
| 4 | Technical Specifications
| 4.1 | Solid-state/microprocessor-controlled frequency generator.
| 4.2 | Monopolar and bipolar outputs, electrically isolated from ground.
| 4.3 | Minimum output frequency is higher than 320kHz.
| 4.4 | Monopolar modes include pure cut, blend, and coagulate (soft, contact and spray).
| 4.5 | Bipolar mode includes coagulate and cut mode.
| 4.6 | Maximum monopolar cut power output maximum 300 W.
| 4.7 | Maximum monopolar coagulation power output maximum 100 W.
| 4.8 | Maximum bipolar power output maximum 100 W.
| 4.9 | Hand switch mode when button-activated probes are connected.
| 4.1 | Foot switch that can operate in monopolar and bipolar modes.
| 4.11 | Yellow buttons/pedals for cut and blue buttons/pedals for coagulate.
| 4.12 | Grounding pad/return electrode monitored for patient connection.
| 4.13 | Front panel allows mode selection, power settings and on/off.
| 4.14 | Display shows output power, system errors and electrode failure.
| 4.15 | Automatic shut off/generator deactivation on grounding pad/return electrode connection failure.
| 4.16 | Audible and visual indicators of activation and alarms.
| 4.17 | Self-test mode.
| 4.18 | Protection against defibrillator discharges.
| 4.19 | The unit shall come with trolley well designed to fit the generator with drawers for keeping the accessories
| 5 | Accessories, Spare Parts and Consumables
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer
| 5.2 | 1 x Foot switch, two pedals, yellow and blue, with connecting cable.
| 5.3 | 2 x Grounding pad/return electrodes, reusable, with 3m connecting cable (adult & child).
| 5.4 | 2 x Monopolar electrode handle, reusable, foot switch controlled, with connecting cable.
| 5.5 | 2 x Monopolar electrode handle, reusable, finger switch controlled, with connecting cable.
| 5.6 | 1 x Set different monopolar reusable electrodes (needle, Knife, ball, and Wire loop).
| 5.7 | 2 x Bipolar forceps, reusable, foot switch controlled, with connecting cable (short, straight, tip-angled).
| 5.8 | 2 x Bipolar forceps, reusable, foot switch controlled, with connecting cable (long, straight, tip-angled).
| 5.9 | 1 x Trolley on 4 antistatic swivel castors, 2 with brakes, fit with a drawer and storage for foot pedal/switch
| 6 | Operating Environment
| 6.1 | The unit shall be capable of being stored continuously in ambient temperature of 10°C -60°C and relative humidity of 30-90%
| 6.2 | The unit shall be capable of operating in ambient temperature of 10°C-40°C and relative humidity of 30-75%
| 7 | Power supply
| 7.1 | Input power supply: 220/240 V AC , 50Hz single phase schuko plug
| 8 | Standards & Safety Requirements
| 8.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 8.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
| 8.3 | IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment
| 9 | User Training:
| 9.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
| 10 | Warranty
| 10.1 | Comprehensive warranty for 2 years.
| 11 | Maintenance Service During Warranty Period
| 11.1 | Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 12 | Documentation
| 11.1 | User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 12.2 | Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 13.3 | List of important spare parts, accessories and consumables with their part numbers and costing.
3 |  | Infant Incubator with UPS
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | An infant incubator provides a closed, controlled environment that warms an infant by circulating heated air over the skin. The heat is then absorbed into the body by tissue conduction and blood convection. Ideally, both the skin and core temperatures must be maintained with only minor variations.
| 2 | Operational Requirements
| 2.1 | High quality with humidity and servo controlled double walled with cabinet incubator. The quality of the material shall be very high and crystal transparent.
| 3 | System Configuration
| 3.1 | Infant Incubator Servo Control, complete unit with all standard accessories.
| 4 | Technical Specifications
| 4.1 | Microprocessor controlled, easy access control panel with feather touch switches. Facility for self-test function.
| 4.2 | Shall have facility to elevate base to offer adjustable range.
| 4.3 | Facility with both servo control as well as air temperature control and servo humidifier.
| 4.4 | Shall have Oxygen port with tubing.
| 4.5 | Continuous bed tilt up to 8° on either sides
| 4.6 | Head end raise facility with auto lock.
Both visual and audible alarms for:
| 4.7 |  High / low temperature.
 Air circulation / probe / system / power failure.
 Humidity control.
| 4.8 | Head end raise facility with auto lock.
| 4.9 | Facility to take x-ray and weight without removing baby.
| 4.1 | Facility to display and trending of temperature information on compatible monitors with other physiological parameters.
| 4.11 | Height 140 +/- 5 cm, depth at least 60 mm, width at least 90 mm.
| 4.12 | Mattress to hood distance (approx.) 40 cm.
| 4.13 | Working level (approx.): 90 to 100 cm.
| 4.14 | Iris port for tubing, leads, probes.
| 4.15 | Gel mattress (approx.) 4 cm. thick and easily washable.
| 4.16 | With at least 4" dia. castors wheel with swivel in all directions and with front lockable wheels. Two shelves cabinets with door.
| 4.17 | Weight (approx.) 90-100 kg.
| 4.18 | Patient control (Servo) mode, 35 oC to 37 oC and Air control (manual mode), 20 oC to 37 oC
| 4.19 | Air velocity less than 10 cm/sec with inner wall.
| 4.20 | Temperature variability less than +/- 0.2 oC. and temperature resolution +/- 1 oC.
| 4.21 | Average oxygen input concentration range 5-15 litres/min or 25-70%.
| 4.22 | Humidification:
 Standard: 10-75% dependent on nursery environment and incubator temperature setting.
 Servo: 40-80% regardless of nursery environment.
| 4.23 | Double wall canopy with six hand ports.
| 4.24 | Shall accommodate IV pole.
| 4.25 | C02 flushing, according to lEC 60601-2-19 / 105.1 Maximum C02 concentration inside incubator 0.2%.
| 4.26 | Servo control for Oxygen with integrated monitoring.
| 4.27 | Noise level < 49 dB.
| 5 | Accessories, spares and consumables
| 5.1 | Accessories:
 Two sets of extra mattresses.
 Two sets of extra sensors.
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 7.3 | Shall meet IEC 60601-2-50 Medical Electrical Equipment PART 2: Particular Requirements for the Safety of Baby Incubators.
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years after acceptance.
| 10 | Maintenance Service during Warranty Period
| 10.1 | During warranty period supplier must ensure preventive maintenance corrective/breakdown maintenance whenever required.
| 11 | Documentation
| 11.1 | User (Operating) manual in English.
| 11.2 | Service (Technical / Maintenance) manual in English.
| 11.3 | List of important spare parts and accessories with their part numbers and costing.
| 11.4 | Certificate of calibration and inspection from factory.
4 |  | Adult Scale With Tape
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | Used for routine height and weight measurements of patients, paediatric to adult.
| 2 | Operational Requirements
| 2.1 | It must be a platform type of weight and height measuring scale on which the patient can stand for measurement of weight and height.
| 3 | System Configuration
| 3.1 | Weighing Machine with Height Measuring Scale, Mechanical , paediatric to adult patients, complete unit.
| 4 | Technical Specifications
It must measure the weight in kilogram.
It must measure the height in centimetre.
Capacity weight: up to 200 kg.
Graduation: ≤ 100 g.
Base Measurement (platform): ≥ 330 x 80 x 340 mm.
It must be mounted on transport castors with breaks to allow free mobility from one place to other.
Required Accessories: Offer must include telescoping measuring rod up to minimum of 200 cm.
| 5 | Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Climate ,Temperature , Humidity, etc. for Sudan.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE or USFDA approved product certificate.
| 8 | Warranty
| 8.1 | Comprehensive warranty for 2 years.
| 9 | Maintenance Service During Warranty Period
| 9.1 | Standard warranty conditions are applicable.
| 10 | Installation and Commissioning
| 10.1 | The supplier must accomplish proper commissioning of the item onsite.
| 11 | Documentation
| 11.1 | User and/or service manual shall be supplied in English.
5 |  | Stethoscope Dual Head
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | To be used for listening to the heart beat of a human, blood flow, bowel sounds and lungs function for adult .
| 2 | Technical specifications
| 2.1 | Dual-head stethoscope for adult use
| 2.2 | Y tube treated rubber with 10mm diameter
| 2.3 | Chest-pieces: stainless steel or chrome brass, with Φ43 mm diaphragm
| 2.4 | Sensitivity of 3.2 dB in a range from 50 to 500 Hz for cardiology
| 2.5 | Sensitivity of 8.1 dB in a range from 600 to 1500 Hz for pneumology.
| 2.6 | Arm: stainless steel or chrome brass, with spring to give lasting spring and maximum reliability and comfort.
| 2.7 | Diaphragm material Epoxy/Fiberglass
| 2.8 | Tubes: to be synthetic material
| 2.9 | Removable plastic ear-pieces
| 3 | Accessories, spares and consumables
| 3.1 | Ear-pieces: 1 pair
| 3.2 | Diaphragm: 3 nos
| 4 | Environmental factors
| 4.1 | The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
| 4.2 | Thu unit shall be capable of operating in ambient temperature of 20 -50 deg C and relative humidity of 80%
| 5 | Standards and safety
| 5.1 | Should be FDA or CE approved product
| 6 | Warranty
| 6.1 | Comprehensive warranty for 2 years.