## RFQ 7-2025
1 |  | Manual Patient Bed One Movement With Mattress
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | A hospital bed is a bed specially designed for hospitalized patients in need of patient ease. These beds have special features both for the comfort and well-being of the patient and for the convenience of hospital staff.
| 1.2 | Mattress is to provide a comfortable platform to rest or sleep upon the bed.
| 2 | Operational Requirements
| 2.1 | The patient bed shall be made of solid steel construction with anti-corrosive and antirust treated baked hard epoxy powder coating
| 3 | System Configuration
| 3.1 | Hospital Bed epoxy powder coated
| 4 | Technical Specifications
| 4.1 | Bed base shall be anti-corrosive and antirust treated epoxy powder coated welded steel bar or epoxy powder coated 18G perforated sheet top to improve ventilation.
| 4.2 | The patient bed shall be fixed height with 2 sections where the backrest section could be elevated by mechanical hand crank located at the foot-end of the bed.
| 4.3 | Shall have 4 IV rod receptacles and mosquito net pole receptacles at the 4 corners
| 4.4 | It shall come with one dual hook anti-corrosive and antirust treated epoxy powder coated or 304 grade stainless steel IV rod.
| 4.5 | Shall have provisions to fix urinary bag on both sides.
| 4.6 | It shall mobile on 4 x 200mm robust swivelling castors with non-marking grey tyres and with at least 2 diagonal castors shall have locking/brake mechanism.
| 4.7 | Bedhead and foot-end panel (head and foot bows) shall be made of either between 4-6 anti-corrosive and antirust treated epoxy powder coated vertical tubular tube of not less than 30mm in diameter or epoxy powder coated steel panel
| 4.8 | Both bedhead and foot-end panel shall be detachable.
| 4.9 | The height of the bedhead panel: not less than 1060mm from floor.
| 4.1 | The height of the foot-end panel: not less than 820mm from floor.
| 4.11 | Overall approximate dimension: not less than 1980mm length, 910mm width, 600mm height
| 4.12 | The mattress negotiable with hospital beds having an overall approximate dimension of not less than 1980mm length, 910mm width.
| 4.13 | It shall be fire retardant antibacterial treated high density approx. 40kg/m3 PU foam mattress.
| 4.14 | The mattress shall have thickness of at least 100mm.
| 4.15 | Mattress with two sections shall be designed to bend with the positioning of the bed when the backrest and foot section of the bed are adjusted.
| 4.16 | The weight capacity of the mattress shall be more than 100kg.
| 4.17 | Mattress Cover:
The mattress shall come with a zipped fire retardant antibacterial, antistatic, acid resistance, waterproof and washable vinyl or vinylized nylon cover. It shall be designed to provide ventilating airflow over the patient’s skin. The zip shall be a heavy-duty/large toothed synthetic zipper to enable inspection of the inner foam and totally covered by a flap extending over the zip to prevent ingress of fluids into the actual mattress and allow for replacement.
| 5 | System Configuration Accessories, spares and consumables
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer. Bidders shall specify, in a separate Excel worksheet, the quantity and details of any items included in this offer which have not been specified in this Technical Specifications Form.
| 6 | Operating Environment
| 6.1 | The system offered shall be designed to store and to operate normally under the conditions. The conditions include Climate, Temperature, Humidity, etc.
| 7 | Warranty
| 7.1 | Warranty for 2 years.