## RFQ 8-2025
RFQ 08-2025- Specs
1 |  | Dental Unit with full accessories
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | A Dental chair for Dental treatment
| 2 | Operational Requirements
| 2.1 | It shall operate on AC power supply.
| 3 | System Configuration
| 3.1 | Dental Unit with complete attachments and accessories.
| 3.1 | Patient chair, 1 unit
| 3.2 | Working stool, 2 units
| 3.3 | High speed hand piece, 2 units
| 3.4 | Low speed motor with 1 straight hand piece & 1 contra-angle hand piece, 1 set
| 3.5 | A built-in light cure unit
| 3.6 | A built-in ultrasonic scaler with one each of pointed and flat scaler tips
| 3.7 | One air compressor at least 1 horse power
| 3.8 | One suction unit at least 1 horse power
| 3.9 | One unit of automatic amalgamator
| 3.1 | One set of amalgam carrier with tips and one amalgam well
| 3.11 | One set of 4 pieces of amalgam condensers
| 3.12 | One set of 5 pieces of amalgam carvers
| 3.13 | One set of 5 pieces of burnishers
| 3.14 | Bidder shall indicate brand and model information here and provide technical data document for major components specified above.
| 4 | Technical Specifications
| 4.1 | Patient type: adult & paediatric & deformity.
| 4.2 | Main unit standard configuration as follow:
| 4.3 | Patient chair:
| 4.4 | With electrical chair movement and deluxe double articulated headrest.
| 4.5 | With one left armrest as standard.
| 4.6 | 3 pre-set chair positions: start, treatment and rinsing position.
| 4.7 | Electrical patient chair loading capacity: not less than 1323N (135kg).
| 4.8 | Backrest movement range 105°-175°.
| 4.9 | The lowest position of the patient chair from the ground shall not be less then 380mm.
| 4.1 | The highest position of the patient chair from the ground shall not be less than 780mm.
| 4.11 | Chair movement is controllable by the 4 way foot control at the chair base without touch panel.
| 4.12 | No cables on the floor, hygienic and clean.
| 4.13 | One main switch to control air, water and power.
| 4.14 | The chair position is locked while an instrument is working.
| 4.15 | With chair-backrest safety system, backrest and seat movement can be stopped once it meets obstacle.
| 4.16 | Dentist element:
| 4.17 | Dentist element with whip arm system.
| 4.18 | Height of dentist element is adjustable.
| 4.19 | 1 X-ray film viewer (12V, 2000cd/m2).
| 4.2 | 1 silicon mat for the dentist element which can be sterilized.
| 4.21 | 1 three way syringe.
| 4.22 | 3 ISO 4-hole/Midwest hand piece hoses.
| 4.23 | 1 air pressure meter.
| 4.24 | Assistant element:
| 4.25 | 1 three way syringe.
| 4.26 | 1 strong suction hose.
| 4.27 | 1 saliva ejector.
| 4.28 | With suction filter system.
| 4.29 | Water unit:
| 4.3 | The cuspidor can be swivelled and removable for easy cleaning.
| 4.31 | Cup filler and bowl rinsing systems shall prevent over filling of cup and prolong rinsing of bowl. Preferably programmable.
| 4.32 | With automatic water heating system (24V)
| 4.33 | With water venturi and air water separator system
| 4.34 | Fresh water bottle, at least 1.5L
| 4.35 | Operating light:
| 4.36 | Colour temperature: 3800-4500K.
| 4.37 | 12V, 50 Watt Halogen bulb or Better.
| 4.38 | Dental light intensity: min 25000lux with intensity dimming function.
| 4.39 | Working stool, 2 units.
| 4.4 | Mobile on 5 castors.
| 4.41 | Height of seat and backrest is adjustable.
| 4.42 | Backrest angle is adjustable and lockable.
| 4.43 | Come with NSK or equivalent high speed hand piece, 2 units.
| 4.44 | Come with NSK or equivalent low speed motor with 1 straight hand piece & 1 contra-angle hand piece, 1 set.
| 4.45 | Come with a built-in light cure unit.
| 4.46 | Come with a built-in ultrasonic scaler with one each of pointed and flat scaler tips.
| 4.47 | Come with one unit of automatic amalgamator.
| 4.48 | Come with one set of amalgam carrier with tips and one amalgam well.
Bidder shall indicate brand and model/ part number information here and provide catalogue of each piece of item offered here.
| 4.49 | Come with one set of 4 pieces of amalgam condensers, one each of serrated small size, serrated big size, smooth small size & smooth big size.
Bidder shall indicate brand and model/ part number information here and provide catalogue of each piece of item offered here.
| 4.5 | Come with one set of 5 pieces of amalgam carvers, double-ended, made of stainless steel, one size each from small, medium up to large size.
Bidder shall indicate brand and model/ part number information here and provide catalogue of each piece of item offered here.
| 4.51 | Come with one set of 5 pieces of burnishers, double-ended, made of stainless steel, one size each of different type.
Bidder shall indicate brand and model/ part number information here and provide catalogue of each piece of item offered here.
| 4.52 | Come with one air compressor at least 1 horse power or capacity sufficient to supply to the Dental chair specified above, whichever higher. Bidder shall indicate capacity of the unit offered here.
Air compressor must be oil less & noise less.
| 4.53 | Come with one suction unit at least 1 horse power or capacity sufficient to supply to the Dental chair specified above, whichever higher. Bidder shall indicate capacity of the unit offered here.
| 5 | Accessories, Spare Parts and Consumables
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
| 6 | Operating Environment
| 6.1 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metres in length.
| 6.2 | Environment:
 Temperature:10-45 degree C
 Relative Humidity: not more than 98%
| 6.3 | Air supply pressure 0.55~0.80Mpa
| 6.4 | Water supply pressure 0.20~0.40Mpa
| 7 | Standards & Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 7.3 | Electrical safety conforms to standards for electrical safety IEC-60601.
| 8 | User Training
| 8.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
| 9 | Warranty
| 9.1 | The warranty period for this item shall be 24 months after acceptance of the Goods
| 10 | Maintenance Service During Warranty Period
| 10.1 | Preventive and corrective maintenance services during warranty period shall be included.
| 11 | Installation and Commissioning
| 11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
| 12 | Documentation
| 12.1 | User (Operating) manual in English
| 12.2 | Service (Technical / Maintenance) manual in English
| 12.3 | List of important spare parts and accessories with their part number and costing.
2 |  | Panorama X-Ray Machine
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Technical Specifications
A. X-Ray Tube
1. Generator: High frequency DC generator, operating frequency at least 40kHz
2. Line Voltage: 230V- 240V, 50-60Hz
3. Tube voltage: 57-85KV
4. Tube Current: 10 mA
5. Exposure compensation: Automatic Exposure Control.
6. Focal spot size: : Between 0.7mm X 0.7 mm and 0.5mm X 0.5 mm as per IEC standards
7. Total Filtration: Equivalent to 2.5mm Al or more
8. Magnification of exposures: Should have a constant magnification
9. Effective exposure time: Adjustable Timer (Pediatric & Adult) for the available imaging programmes.
B. Basic Unit
1. The x ray machine should have AERB approval certification.
2. Should be floor or wall mounted model.
3. Should have fully/semi motorized patient positioning system
4. Should have 3 laser positioning lights – Mid Sagital, Frankfurt and Canine
5. Patient positioning by Bite block, Chin rest, Chin support and motorized/manual temple support.
6. Should have accessibility for wheel chair patients.
7. Should have centralized control system with all functions controlled by colour TFT touch screen or PC, Graphical User Interface (GUI) with digital display of technical factors and selected programmes.
8. Should have an adjustable form of focal trough depending on jaw shape and size of patient.
9. Should have automatic compensation for the cervical vertebrae shadow.
10. The imaging geometry should eliminate the redundant shadows and ghost images.
11. Should have a test mode which disables X-ray radiation during operation.
12. Should have exposure counters separate /or combined for Panoramic, Cephalometry and total no of images
13. Should have the following programs:
a) Standard Panoramic Program.
b) Pediatric program (low dose, reduced exposure area for small patient)
c) Automatic Double TMJ Program
d) Sinus program
e) Lateral and PA Cephalogram.
f) Should Support carpus / hand wrist imaging.
g) Submentovertex views.
14. Should have motorized up and down movement.
C. Digital Cephalometric System
1. Should have Computerized automatic cephalometric system.
2. Should have automatic alignment of radiation source
3. Should have functionally designed and easy-to-use head positioner.
4. Should have swiveling nasal support.
5. Should have magnification scale that appear on the image
6. Automatic soft-tissue filter based on the position of the nasion support.
D. Digital Sensor System
1. Should have a high resolution CCD /CMOS flat panel sensor for panoramic and cephalometry. It should be possible to take both panoramic and cephlometric image by using the sensor provided. The cost of one additional sensor shall be quoted separately in the BOQ and the cost offered will not be taken for evaluation.
2. CCD pixel size: between 35?m and 150 ?m.
3. CCD active surface: to be specified.
4. Image field: to be specified.
5. Pixel matrix: to be specified.
6. Soft tissue filter should be software operated.
E. PC Configuration Requirements
Latest branded computer with following specifications
1. Processor –core i5 4th generation or better or equivalent
2. 8 GB RAM and 1 TB or more on Hard disk.
3. Should include 21” LCD/TFT monitor with high resolution to ensure image quality.( Display /monitor resolution min 1024 x1024 or higher)
4. CD/DVD Burner.
5. Graphics card NVIDIA 2 GB dedicated GPU or equivalent.
6. Accept images from CCD/CMOS sensor without any loss of data.
7. Capable of archiving & printing selected image to a standard DICOM destination in DICOM 3.0 format (or newer versions).
8. Storing images in the local disk for predefined period.
9. Should be with upgradeable imaging software.
F. Functional Requirements For Work Station
1. Built in routine for using predefined image processing parameters for image quality enhancement.
2. Mechanisms for storing the patient image based on name, date, etc.
3. Capability of storing user defined image processing parameters.
4. Capability of overwriting predefined image parameter with user defined parameters & storing these two images separately.
5. Correcting typographically in patient demographic module, in case the RIS connection was down and manually data entry was done.
6. Capability of changing W/I, flipping, rotating, zooming, collimating annotating incoming image.
7. Auto routing incoming image to predefined DICOM store (SCP storage) or Print destination (SCP print destination).
8. Mechanism of printing multiple images in one film, with the possibility of slide and True size printing.
9. The work station should let the user select a film size and lay out for hard copies with a WYSIWYG (What you see is what you get), output look up table selection, number of copies selection capabilities.
10. Minimum of 15 pre configured layouts for printing and also possibility of further configuration should be possible.
11. Should be provided with table and chair for workstation.
12. Should provide necessary cabling work for additional workstations.
13. Operating system: Windows 7 or more
G. Software specifications
The system should include the following SW applications as standard:
1. Special attention should be placed on pediatric applications.
2. Software for printing on any DICOM printer.
3. Software for storing images on any DICOM3 or newer versions compliance stations.
4. Software for e-filming application: Any Software upgradeability for 5 years to be done by the company
H. Dry Imager System
1. Print images from work stations.
2. Should be capable of printing images in DICOM format / Direct DICOM Compatibility.
3. Mechanism to print images of OPG, Ceph and multiple image prints of IOPA. The tray should be adjustable to (8" x 10") and (10" x 12").
4. Resolution should be 300- 508 dpi or better.
I. Accessories & consumables
1. Films for Dry Imager system – (8" x 10")- 300 numbers
2. Three folded leaded protective barrier – 1 no.
3. 5 KVA Online UPS with 30 minutes backup for operating all the above mentioned equipments.
4. Lead apron light weight Velcro type 1 no, Thyroid guard -1 No, AERB Approved
| 2 | Accessories, spares and consumables
| 2.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer.
| 3 | Operating Environment
| 3.1 | The unit shall be capable of operating in ambient temperature of 10°C-45°C and relative humidity of 10-95%
| 3.2 | The unit shall be capable of being stored continuously in ambient temperature of -20°C -60°C and relative humidity of 10-95%
| 4 | Power supply:
| 4.1 | Input power supply: 220/240 V AC , 50Hz single phase schuko plug
| 5 | Standards and Safety Requirements
| 5.1 | Must submit ISO 9001 or ISO 13485:2003/AC: 2007
| 5.2 | CE or USFDA or TUV approved product certificate.
| 6 | Training
| 6.1 | Must provide user & service training.
| 7 | Warranty
| 7.1 | Comprehensive warranty for 2 years after acceptance.
| 8 | Maintenance Service During Warranty Period
| 8.1 | Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 9 | Documentation
| 9.1 | User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 9.2 | Service (Technical / Maintenance) manual in English Should provide 2 sets (hardcopy and soft-copy)
| 9.3 | List of important spare parts, accessories and consumables with their part numbers and costing.
3 |  | Echo Ultrasound Machine with probes for Adult and Pediatric and with printer and UPS
| No. | Item Specifications | Fill your Specification
| 1 | Description of Function
| 1.1 | Colour Doppler Echocardiography System is required to study the anatomical abnormalities and blood flow in the heart and associated vessels.
| 2 | Operational Requirements
| 2.1 |  Latest generation Electronic Phased array Colour Doppler system with minimum 512 electronic independent channels.
 System must be DICOM ready and capable of being interfaced with HIS/RIS/PACS.
 Must be upgradable to next generation system on site.
 Frequency compounding or better technology for better resolution and penetration.
| 3 | System Configuration
| 3.1 | Colour Doppler System with all application packages, quad loop for serial studies with high frame rate review. Harmonic imaging capability in all modes. (Tissue, Contrast, Anglo) Integrated Stress Echo Package, Digital Storage and Retrieval – 01no.
 1-3 MHz Adult Cardiac probe Electronics Phased Array probe- 01 ea.
 3-11 MHz Electronics Phased Array Probe for Vascular applications- 01 ea.
 Multi-plane TEE Probe: 4-8 MHz for Adult as well as Paediatric echocardiography.
 5-10 MHz Electronic phased array probe for Paediatric cardiology.
 Colour Printer -01no.
 B/W Video Thermal Printer -01no.
| 4 | Technical Specifications
| 4.1 | Latest generation Electronic Phased array colour Doppler system with Minimum 512 Electronic independent channels.
| 4.2 | 256 gray shades for sharp contrast resolutions
| 4.3 | Multi-dimensional Beam former for generating two images simultaneously-one at low end of bandwidth and one at high end-then selectively retrieves and mixes the components together for finely textured 2-D or B mode image with superior contrast resolution
| 4.4 | Adult Cardiac and Vascular Probes to be supplied which must be latest generation wide band transducers without frequency selection for higher sensitivity of response over a broader frequency range of operation. All probes to be phased array.
Probes for paediatric application and Trans oesophageal Echo for future requirement.
| 4.5 | Harmonic Imaging: System must have following modes in harmonic with separate setting for:
 Tissue Harmonic.
 Contrast Harmonic - both triggered and real time
 Harmonic Angio.
 Quantification of harmonics imaging
| 4.6 | Harmonic imaging capability in Adult Cardiac, Paediatric Cardiac and linear Probe.
| 4.7 | Gain control in two dimensions for additional level of flexibility to image quality control.
| 4.8 | Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes
| 4.9 | Frame rate must be 300 FPS or more.
| 4.1 | Steerable PW/CW in all Phased Array probes.
| 4.11 | High definition acoustic zoom for enlarging sections of 2D and colour flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate.
| 4.12 | Modes - 4D, M-Mode, Steerable PW/CW Doppler, Colour Doppler, and High Definition colour flow with capability of automatically picking up colour flow as a function of focal depth
| 4.13 | Monitor must be 15" or more, high resolution colour monitor.
Tilt and Swivel monitor must be able to view in all angles and all light conditions.
| 4.14 | Colour Flow Imaging for:
 Increased lateral & spatial resolution.
 Detection of even subtle areas of turbulence, displaying a more physiological blood flow appearance without loss of frame rate.
 Colour flow with capability of automatically picking up colour flow as a function of focal depth
| 4.15 | Tissue Colorization (B-colour) for improved contrast resolution
| 4.16 | Application software for Adult, Paediatric, Foetal and Peripheral Vascular and Trans oesophageal applications. (All application packages must be built into the system).
| 4.17 | Cine loop memory- more than 120MB of memory.
 High Frame rate review for better clarity of playback images study in slow motion.
 Quad loop with memory for pre and post image comparison of any procedure.
 Memory: 256 frames or more in quad loop. M Mode & Doppler Scroll Memory-40 seconds or more.
 Frame grabber facility for post analysis.
| 4.18 | Various maps for pre and post processing.
| 4.19 | ECG triggers facility.
| 4.2 | User defined system and application pre-sets for multi-user department.
| 4.21 | Minimum 4.8 GB optical disc drive for image storage and retrieval. (standard with system)
| 4.22 | Dedicated integrated dynamic stress echo package for flexible user defined protocols with stacked sub loops facility and contrast stress protocol.
| 4.23 | Tissue movement colorization with quantification possibility for IHD/CAD patients.
| 4.24 | Three transducer ports will be preferred.
| 4.25 | Colour Map resolution up to 128 levels.
| 4.26 | Study Manager (> 1.5 GB) for on-cart digital acquisition, review and editing of complete patient studies.
| 4.27 | Facility of Real time perfusion studies
| 4.28 | System Peripherals shall include:
 CD Writer with calculation facility on playback.
 Colour Video Printer.
 B/W Thermal Printer.
| 5 | Accessories, spares and consumables
| 5.1 | Accessories:
 DVD/CD Recorder with 100 CDs and 100 DVDs
 Colour Print Paper- 500 sheets
 B/W Thermal Paper - 10 rolls
 ECG Cable - 02nos.
 MO Disc - 10pcs
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
| 6.3 | UPS of suitable rating conforming to international standards shall be supplied for minimum 30 min. backup for the entire system.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 7.3 | The product shall comply to IEC 60601-2-37 ed1: Medical Electrical Equipment - Part 2-37: Particular Requirements for the Safety of Ultrasonic Medical Diagnostic and Monitoring Equipment.
| 7.4 | Type of protection against electric shocks - Class I
Degree of protection against electric shocks for ultrasound probes Type "BF"
For ECG electrodes Type 'CF"
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years.
| 10 | Maintenance Service during Warranty Period
| 10.1 | During the warranty period supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 11 | Documentation
| 11.1 | User (Operating) manual in English
| 11.2 | Service (Technical / Maintenance) manual in English
| 11.3 | List of important spare parts and accessories with their part numbers and costing.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
4 |  | Stress ECG Machine Complete with Treadmill and with software with UPS
| No. | Item Specifications | Fill your Specification
1. Description of functions:
In this system, the patient exercises on a treadmill according to a standardized protocol and the cardiac abnormalities can be studied under stress conditions which we may miss under resting
2. Operational requirements
2.1 The treadmill stress test system should be complete with acquisition of resting and stress ECG, Treadmill unit with interface with all the protocols and provision of printing the resting as well as stress ECG and analyzing the same
2.2 Should be able to be interfaced to hospital information systems / LAN/WLAN
3. Technical specifications
1. System should have telemetry module (Blue tooth based) between patient and machine
2. USB connectivity for high speed data transfer from acquisition device to PC.
3. Should have integrated automatic programmable stress test NIBP measurement module with Bi-directional data transfer (compatible with the stress test system)
4. Repairable and generic patient cables and TMT/ECG accessories should be used
5. Should have software inbuilt online support system for controlling system through internet for providing online support.
6. Company should make spares available for the entire life of the system.
7. Service response time should be quick and service personnel should be able to attend the call within 48 hours (working days).
8. Quality certifications like ISO-9001:2008, ISO-13485:2012 or equivalent
9. Should have certification from USFDA and/or valid CE certificate from notified body.
Specification for Acquisition Software
1. Should acquire data from 12 Lead simultaneously
2. Should have frequency response from 0.05 – 150 Hz
3. Should have notch filter around 50 Hz
4. Should have sampling rate of at least 4500 samples/Second
5. Should have rejection ratio (CMRR) more than 100 dB
6. Should have internal Defib protection
7. Should have less than 10 uAmp patient leakage
8. Should have input impedance of more than 100 MOhms
Specification for stress test software
1. Should be capable of displaying real-time or stored ECG tracings.
2. Should display and regularly update ECG 12 leads, 12 medians, 1 expanded median, HR, BP, METS, Stage time, test time, protocol name, stage name, speed and grade of treadmill.
3. Should be able to give standard ECG report including standard ECG analysis.
4. Should have automatic stage print out facility at the end of each exercise
5. Median complexes of all 12 leads with rhythm leads with different formats should be available
6. Should have capability to display real-time ST running trend
7. Should have ability to display trend graph for HR, BP, ST level, ST slope and J amplitude
8. Should have automatic detection, display, storage and review of rhythm events
9. Should be able to display ECG in various formats like 3 Lead + 12 Median; 6 Lead + 12 Median; 12 Lead + 12 Median
10. Should have controllable display sweep speeds ranging from 5.0 to 100 mm/Sec
11. Should have display sensitivities ranging from 0.25 to 8.0 cm/mV
12. Should have base line correction (BLC) for stable baseline during test
13. Should run various test protocols like Bruce, Modified Bruce, Balke, Ellested, Naughton and user defined protocols.
14. Acquisition and analysis softwares should be upgradable to latest version free of cost.
15. Should have capability to import patient data from HIS and also manually edit/add data of patient
16. Raw data from software should be made available in standard formats for further analysis with softwares like MATLAB.
17. Should be able to print report in PDF format.
18. Should able to print reports with standard Laser printers on A4 Plain sheets.
Specifications for treadmill
Treadmill should have:
1. Fully interfaced and controllable from software
2. Controllable speed of 9-20 Km/H
3. Variable inclination (grade) from 0 – 22%
4. Adequate walking area ~ 1500 mm X 500 mm
5. Controlled by optically isolated RS 232 or USB
6. AC motor capable of running on inverter/UPS
7. Emergency stop feature
8. Power requirement 230 V,50 Hz, 15 A
Specification for computer
System should be supplied with Trolly mounted branded all in one PC which should have minimum:
1. i5 processor, 16 GB RAM, 2TB Hard Disk, DVD/CD Drive, ~ 23” touch screen LED display.
2. Windows 11 (64 bit) Operating System.
3. Laser Printer for printing on A4 Sheets.
4. 6 KVA online UPS for entire system backup for at least one hour.
Should be supplied with required standard accessories:
1. Disposable standard chest electrodes - 100 Pcs
2. Reusable ECG electrodes – 2 sets
3. Extra Patient cable - 01
4. Spike protector (with 5 sockets) - 01 No
5. Earthing cord – 1 No
6. L-N Key set
7. ECG jelly – 2 No
8. A4 Sheet Ream – 1 No.
9. Fuse (1.5 Amp and 10 Amp) – 10 each or compatible
10. User Manual – 01 No.
11. Cuff motion sensor – 01 No.
12. NIBP Module – 01 No.
13. Wrist strap of NIBP cable – 01 No.
14.Must submit ISO13485:2003/AC:2007 for Medical Devices ANDCE or US FDA approved product certificate.
15.Comprehensive warranty for 2 years after acceptance.
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
5 |  | Auto Hematology Analyzer 3 Part diff. with Reagent set with UPS
| No. | Item Specifications | Fill your Specification
| 1 | Description of Function
Automated haematology analyser or complete blood cell counter is used to count various types of blood cells in the blood.
| 2 | Operational Requirements
Fully automated 3 parts differential haematology analyser.
| 3 | System Configuration
Fully Automated Haematology Analyser, complete unit with all standard reagents, consumables and
accessories.
| 4 | Technical Specifications
Determination of 18 to 19 parameters, with 3-part differential, for routine haematology.
Shall have fully automatic, open system.
Sample volume: < 30ul.
Throughput: approx. 50 samples per hour, 24h power on, with dormancy and wake function.
Determination of: Red blood cell(RBC), White blood cell(WBC), Haemoglobin(HGB), Haematocrit(HCT), Mean cell volume(MCV), Mean cell haemoglobin(MCH), Red cell distribution(RDW-SD and RDW-CV), Platelets(PLT),Platelet distribution(PDW-SD and PDW-CV), Mean platelet volume(MPV), differential leucocytes (LYM, LYM%, MID, MID%, GRA, GRA%).
Method: Photometry and impedance technology, cyanide-free colorimetry for haemoglobin counting.
Calibration: independent automated calibration and manual calibration for minimum two test modes.
Typical counting time: approximately 60 seconds for differential.
Shall have with self-test capability.
Display: LCD screen.
Indication of self-test failures and assistance messages, sample ID, date and time are reported with test results.
Supplied complete with dedicated data analysis and data management software.
Results are reported on external laser printer.
Shall have built-in RS232, USB2.0 or equivalent, for allowing data transfer and network capability via LIS.
On board memory for about 100-150 tests records.
Shall quote rates for reagents & consumables, calibrators & controls, printer paper, separately and it must be valid for at least 3 years.
| 5 | Accessories, spares and consumables
Reagents & consumables, calibrators & controls, printer paper to be supplied for 1000 samples.
Shall provide compatible laser printer, 1 no.
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
| 6 | Operating Environment
The product offered shall be designed to be stored and to operate normally under the conditions of Sudan. The conditions include Power Supply, Climate, Temperature, Humidity, etc.
Power supply: 220-240VAC, 50Hz fitted with appropriate plug type D (3 pins). The power cable must be minimum 3 meters long.
Suitable UPS with maintenance free batteries, voltage regulation and spike protection for minimum 30 min. back-up for the entire system including computer and printer shall be supplied with the system.
Standards and Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA approved product certificate.
Shall meet IEC 61010-2-081 safety requirements for electrical equipment for measurement, control, and laboratory use - Part 2-081: Particular requirements for automatic and semi-automatic laboratory equipment for analysis and other purposes.
| 7 | User Training
Must provide user training (including how to use and maintain the equipment).
| 8 | Warranty
Comprehensive warranty for 2 years after acceptance.
| 9 | Maintenance Service During Warranty Period
During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required.
| 10 | Installation and Commissioning
The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
| 11 | Documentation
User (Operating) manual in English.
Service (Technical / Maintenance) manual in English.
List of important spare parts and accessories with their part numbers and costing.
Certificate of calibration and inspection from factory.
6 |  | Colorimeter
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | General purpose colorimeter use in clinical laboratory.
| 2 | Operational Requirements
| 2.1 | Microprocessor controlled system.
| 3 | System Configuration
| 3.1 | Colorimeter with complete accessories.
| 4 | Technical Specifications
| 4.1 | Must have 8 no of filters wave length from 340 nm to 730 nm.
| 4.2 | Must have a 2 digit LED display calibrated directly in optical density.
| 4.3 | Detector must be encased spill proof photocell.
| 4.4 | Must have facilities for concentration, calculation, percentage transmission and optical density.
| 4.5 | Lamp source: Broad spectrum LED or halogen covering full visible range
| 5 | Accessories, spares and consumables
 Square and round cuvette minimum volume 1ml.
 Cuvettes: 10 nos.
 Lamp: 02 nos
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 7.3 | Must comply with IEC 61010 Safety requirements for electrical equipment for measurement, control, and laboratory use
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years after acceptance.
| 10 | Maintenance Service During Warranty Period
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 11 | Installation and Commissioning
| 11.1 | Supplier must accomplish proper installation and commissioning of the equipment on site.
| 12 | Documentation
| 12.1 | User (Operating) manual in English
| 12.2 | Service (Technical / Maintenance) manual in English
| 12.3 | List of important spare parts and accessories with their part numbers and costing.
7 |  | ECG Machine
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | Used to detect the electrical signals associated with cardiac activity and produce an ECG, a graphic record of the voltage versus time.
| 2 | Operational Requirements
| 2.1 | Microprocessor controlled digital 3 channel ECG machine suitable for adult, paediatric and neonate applications.
| 3 | System Configuration
| 3.1 | Three channel ECG machine with complete accessories.
| 4 | Technical Specifications
| 4.1 | Simultaneous 12 lead acquisition: aVR, aVL, aVF, I, II, III and V1-V6 derived from 10 electrodes (RA, LA, RL, LL, V1 - V6).
| 4.2 | Arrythmia and ST elevation detection.
| 4.3 | Minimum gain/sensitivity settings include 2.5, 5, 10 and 20 mm/mV.
| 4.4 | Adjustable trace speeds include 5, 10 (and/or 12.5), 25 and 50mm/sec.
| 4.5 | Minimum HR range 30 – 300 bpm with rhythm analysis.
| 4.6 | Minimum guaranteed diagnostic frequency response of 0.05 – 150 Hz.
| 4.7 | Common Mode Rejection Ratio (CMRR) at 60 Hz > 105 dB or better.
| 4.8 | Calibration signal of 1 mV, manual and/or automatic.
| 4.9 | Selectable/adjustable filters for baseline drift, muscle artefacts, mains power.
| 4.1 | Pacemaker detection.
| 4.11 | Accuracy of input signal reproduction ± 5 % or ± 40 μV, whichever is greater.
| 4.12 | Input impedance > 50 MΩ.
| 4.13 | Internal noise level < 12.5 µV peak-to-peak.
| 4.14 | Automatic baseline centring.
| 4.15 | Defibrillation fluctuation/overload protection.
| 4.16 | Baseline recovery < 5 s after defibrillation.
| 4.17 | AC fluctuation protection.
| 4.18 | RECORDER AND PRINTER
Minimum of 3 recording channels.
Recorder display includes date/time, patient data and heart rate and basic settings.
Capable of displaying one group of at least three channels simultaneously.
Recorder waveform display includes lead marker and timing marker.
Integrated/built-in printer.
Capable of printing user selected number of channels.
Capable of printing one group of at least three channels simultaneously.
Paper speeds include 5, 10 (or 12.5), 25 and 50 mm/s.
Compatible with Z-fold paper and optionally with roll paper also (indicate compatibility).
| 4.19 | DATA INPUT/OUTPUT, STORAGE AND ALARMS
Integrated alpha-numeric keyboard.
Patient data input fields include name, age, height and weight, gender.
Backlit Liquid Crystal Display (LCD) display screen, minimum of 6 inche diagonally.
Capable of internally storing a minimum of 10 waveforms for later retrieval, printing and/or transmission.
Expandable storage for additional waveforms if required, via USB.
Capable of exporting waveform data and reports, via USB or LAN.
Visual alarms for patient connection (lead faults), heart rate, printer and paper errors, and system errors.
Automatic self-test at start up.
| 4.2 | Built-in rechargeable lithium-ion battery.
| 4.21 | Minimum battery operating time is 100 ECG exams or 4 hours of continuous recording.
| 4.22 | Automatic switch to battery in case of power failure, automatic recharge on connection to mains.
| 4.23 | Maximum battery charging time to full charge is 8 hours.
| 5 | Accessories, spares and consumables
| 5.1 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools.
| 5.2 | Accessories:
 Reusable Patient cable with reusable electrodes for adult 1 set.
 Reusable Patient cable with reusable electrodes for paediatric 1 set.
 Reusable patient cable with reusable electrodes for neonate & infant- 1 set.
 Extremity clamp electrodes, reusable- 4 nos.
 Recording paper 2 Z-folded or rolls- 5 rolls
 Bottles of electrode gel, approximately 350ml- 1 nos.
 Spare rechargeable battery pack- 1 no.
 Set of spare fuses- 1 set
 Plastic protective dustcover- 1 no.
| 6 | Operating Environment
| 6.1 | The unit shall be capable of operating in ambient temperature of 10°C-45°C and relative humidity of 10-95%
| 6.2 | The unit shall be capable of being stored continuously in ambient temperature of -20°C -60°C and relative humidity of 10-95%
| 7 | Power supply:
| 7.1 | Input power supply: 220/240 V AC , 50Hz single phase schuko plug
| 8 | Standards and Safety Requirements
| 8.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 8.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 8.3 | Electrical safety conforms to standards for electrical safety IEC-60601-1 General Requirements and IEC-60601-2-25 Safety of Electrocardiograms.
| 9 | Training
| 9.1 | Must provide user & service training.
| 10 | Warranty
| 10.1 | Comprehensive warranty for 2 years after acceptance.
| 11 | Maintenance Service During Warranty Period
| 11.1 | Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 12 | Documentation
| 12.1 | User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 12.2 | Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 12.3 | List of important spare parts, accessories and consumables with their part numbers and costing.
8 |  | Delivery Table
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | Delivery bed is used for Baby Delivery and must incorporate ideal blend of the patient's individual requirements on comfort and the professional needs of the delivery team, focusing on the aesthetic and functional design of the entire product.
| 2 | Operational Requirements
| 2.1 | Manually operated delivery bed.
| 3 | System Configuration
| 3.1 | Delivery Bed with complete attachments and accessories.
| 4 | Technical Specifications
| 4.1 | It must have manual adjustments for height and back positions.
| 4.2 | It must have collapsible side rails.
| 4.3 | It must have three sectional mattresses and seat section must have large perennial cut.
| 4.4 | It must have head board which can be detached.
| 4.5 | Must have wheels provided with locking system.
| 4.6 | Must have retractable foot section so as to convert bed into table.
| 4.7 | Must have infusion rods which have adjustable heights, quick release and attaches to all corners of bed.
| 4.8 | Must have adjustable leg rests.
| 4.9 | Must have push grip handles.
| 4.10 | Must have sliding stainless steel bowl at perennial part of table.
| 4.11 | It must have catheter bag holder which can be attached on either side of bed.
| 4.12 | It must be able to give trendelenburg, reverse trendelenburg and 70 degree sitting position.
| 4.13 | It must have adjustable foot supports.
| 4.14 | It must be easy to maintain clean and sterilize (especially blood stains).
| 4.15 | Frame must be (washable) stainless steel.
| 4.16 | Dimensions (approx.):
 Length: 7 feet
 Width: 3 feet
| 4.16 | Capacity load of 180 kg or more
| 5 | Accessories, spares and consumables
| 5.1 | All standard attachments and accessories: 01 set
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature , Humidity , etc. for Sudan.
| 7 | Warranty
| 7.1 | Comprehensive warranty for 2 years after acceptance.
| 8 | Maintenance Service During Warranty Period
| 8.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
9 |  | Minor Surgical Set
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
Technical Specifications
surgical knife handle - Qty. 1
needle holder straight (mayo) 17.5 - Qty. 1
straight artery forceps ,pean 16 - Qty. 1
curved artery forceps, pean 16 - Qty. 1
dissecting forceps - Qty. 1
Curved. Scissors , blunt point ( mayo) 17 cm - Qty. 1
surgical scissors , straight , blunt 14.5 cm - Qty. 1
sharp uterine curette 26 cm x 9 mm sims - Qty. 1
blunt uterine curette 26 cm x 8 mm ( sims) - Qty. 1
kidney basin 475 ml (16oz) stainless steal 12'' x5 x 2-3/8 - Qty. 1
torch power supply 220v , 50Hz - Qty. 1
Bowl, solution, stainless steel, 8 liters. - Qty. 1
gauze scissors - Qty. 1
dressing tray - Qty. 1
surgical towels - Qty. 1
handing forceps - Qty. 1
Operating Environment
The unit shall be capable of operating in ambient temperature of 15°C-35°C and relative humidity of 10-95%
The unit shall be capable of being stored continuously in ambient temperature of -5°C -40°C and relative humidity of 10-95%
Standards & Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
Warranty
Comprehensive warranty for 2 years.
10 |  | Mobile Operating Lamp
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | Mobile operating light is required for carrying out operations in an emergency environment and the system can be moved from place to place in hospitals.
| 2 | Operational Requirements
| 2.1 | Shall operate on mains electric supply as well as on battery.
| 3 | System Configuration
| 3.1 | Operating Light Mobile with single light head, moveable on casters and with all standard accessories.
| 4 | Technical Specifications
| 4.1 | The light shall be designed with good counterbalance mechanism in order to ensure stability of light head in all positions
| 4.2 | Shall have single light head.
| 4.3 | Light head shall not be greater than 400mm diameter
| 4.4 | Number of bulb: 1 main bulb halogen 24V 150W with 1 backup bulb with auto-switching or multi-bulbs light head not more than 3 bulbs of halogen 12V 50W.
| 4.5 | Light intensity range, shall not be less than 80,000 lux at 1 meter distance from light source. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF.
| 4.6 | Light temperature, between 4000 - 4500 K.
| 4.7 | Shall have colour rendering index in between 92 - 94.
| 4.8 | Shall have 99% heat filtrating.
| 4.9 | Temperature increase at head shall not be more than 2 oC.
| 4.10 | Temperature increase on operating field shall not be more than 15 oC.
| 4.11 | Radiant heat energy (beam temperature) shall not exceed 25,000 microwatts per square centimetre measured 1m from the light source.
| 4.12 | The light offered shall have safety designed to prevent patient from burns, especially during the ophthalmic procedures. The light offered shall be certified safe to be used under ophthalmic procedures.
| 4.13 | Working distance range (focal length): 70 - 130cm.
| 4.14 | Depth of field with focused light: > 60cm.
| 4.15 | Luminance field size: 15 ~ 25cm diameter, adjustable.
| 4.16 | Shall have a control to regulate light intensity and to switch on the unit.
| 4.17 | Shall have an On/Off switch at lamp head.
| 4.18 | Shall come with continuous dimmer, continuous focus adjustment, continuous field adjustment.
| 4.19 | Sterilizeable handle to regulate light field size, 2 pcs.
| 4.20 | Vertical adjustment shall not be less than 115cm.
| 4.21 | Rotation 360 degree.
| 4.22 | Come with Ophthalmic procedures safe light bulbs with a minimum of 1000 hours lifespan.
| 4.23 | Interference suppressed complies with VDE 0875 or equivalent.
| 4.24 | Transformer and operating elements shall be integrated in light head housing.
| 4.25 | Mobile Stand:
Shall be based on light weight easily moveable stable support with at least 4 casters with locking counter balance mechanism in order to ensure stability of light head in all positions and with swivel arm. Caster must be medical chemical resistant.
| 4.26 | Battery:
 Lithium ion built in rechargeable batteries with capacity sufficient for operating in battery mode (fully charged) for minimum of 3 hours.
 Shall include a built-in automatic battery charger with proper protection against battery damage.
 Shall include battery power (charge) indicator.
| 5 | Accessories, spares and consumables
| 5.1 | Accessories:
 1 x spare halogen bulb.
 1 x spare set of fuses.
| 5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
| 7.3 | Electrical safety conforms to standards for electrical safety IEC 60601-1 General requirement for Electrical safety of Medical Equipment.
| 8 | User Training
| 8.1 | Must provide user training (including how to use and maintain the equipment).
| 9 | Warranty
| 9.1 | Comprehensive warranty for 1 year after acceptance.
| 10 | Maintenance Service During Warranty Period
| 10.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 11 | Documentation
| 11.1 | User (Operating) manual in English.
| 11.2 | Service (Technical / Maintenance) manual in English.
| 11.3 | List of important spare parts and accessories with their part numbers and costing.
11 |  | Diathermy Machine with Trolley
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of function
| 1.1 | Diathermy Machines are used for surgical cutting and for controlling bleeding by causing coagulation (haemostasis) at the surgical site
| 2 | Operational Requirements
| 2.1 | It shall operate on AC power supply in the operating theatre.
| 3 | System Configurations
| 3.1 | Diathermy Machine (Electrosurgical unit) 300W with complete accessories.
| 4 | Technical Specifications
| 4.1 | Solid-state/microprocessor-controlled frequency generator.
| 4.2 | Monopolar and bipolar outputs, electrically isolated from ground.
| 4.3 | Minimum output frequency is higher than 320kHz.
| 4.4 | Monopolar modes include pure cut, blend, and coagulate (soft, contact and spray).
| 4.5 | Bipolar mode includes coagulate and cut mode.
| 4.6 | Maximum monopolar cut power output maximum 300 W.
| 4.7 | Maximum monopolar coagulation power output maximum 100 W.
| 4.8 | Maximum bipolar power output maximum 100 W.
| 4.9 | Hand switch mode when button-activated probes are connected.
| 4.1 | Foot switch that can operate in monopolar and bipolar modes.
| 4.11 | Yellow buttons/pedals for cut and blue buttons/pedals for coagulate.
| 4.12 | Grounding pad/return electrode monitored for patient connection.
| 4.13 | Front panel allows mode selection, power settings and on/off.
| 4.14 | Display shows output power, system errors and electrode failure.
| 4.15 | Automatic shut off/generator deactivation on grounding pad/return electrode connection failure.
| 4.16 | Audible and visual indicators of activation and alarms.
| 4.17 | Self-test mode.
| 4.18 | Protection against defibrillator discharges.
| 4.19 | The unit shall come with trolley well designed to fit the generator with drawers for keeping the accessories
| 5 | Accessories, Spare Parts and Consumables
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer
| 5.2 | 1 x Foot switch, two pedals, yellow and blue, with connecting cable.
| 5.3 | 2 x Grounding pad/return electrodes, reusable, with 3m connecting cable (adult & child).
| 5.4 | 2 x Monopolar electrode handle, reusable, foot switch controlled, with connecting cable.
| 5.5 | 2 x Monopolar electrode handle, reusable, finger switch controlled, with connecting cable.
| 5.6 | 1 x Set different monopolar reusable electrodes (needle, Knife, ball, and Wire loop).
| 5.7 | 2 x Bipolar forceps, reusable, foot switch controlled, with connecting cable (short, straight, tip-angled).
| 5.8 | 2 x Bipolar forceps, reusable, foot switch controlled, with connecting cable (long, straight, tip-angled).
| 5.9 | 1 x Trolley on 4 antistatic swivel castors, 2 with brakes, fit with a drawer and storage for foot pedal/switch
| 6 | Operating Environment
| 6.1 | The unit shall be capable of being stored continuously in ambient temperature of 10°C -60°C and relative humidity of 30-90%
| 6.2 | The unit shall be capable of operating in ambient temperature of 10°C-40°C and relative humidity of 30-75%
| 7 | Power supply
| 7.1 | Input power supply: 220/240 V AC , 50Hz single phase schuko plug
| 8 | Standards & Safety Requirements
| 8.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 8.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
| 8.3 | IEC 60601-2-2 Medical Electrical Equipment – Part 2-2: Particular requirements for the safety of high frequency surgical equipment
| 9 | User Training:
| 9.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
| 10 | Warranty
| 10.1 | Comprehensive warranty for 2 years.
| 11 | Maintenance Service During Warranty Period
| 11.1 | Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
| 12 | Documentation
| 11.1 | User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 12.2 | Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
| 13.3 | List of important spare parts, accessories and consumables with their part numbers and costing.
12 |  | General Orthopedic Set
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Operational Requirements
| 1 | SS TRAY - BIG (12INCH*15INCH) -Qty.4
| 2 | SS BOWL -BIG -Qty.12
| 3 | KIDNEY TRAY - LARGE -Qty.8
| 4 | KIDNEY TRAY - MEDIUM -Qty.4
| 5 | TOWEL CLIP -Qty.40
| 6 | SPONGE HOLDER -Qty.8
| 7 | BP HANDLE NO. 3 -Qty.4
| 8 | BP HANDLE NO. 4 -Qty.4
| 9 | BP HANDLE NO. 7 -Qty.4
| 10 | LENGAN BAG RETRACTOR - SMALL -Qty.8
| 11 | LENGAN BAG RETRACTOR - MEDIUM -Qty.8
| 12 | LENGAN BAG RETRACTOR - LARGE -Qty.8
| 13 | C-ZERNEY RETRACTOR -Qty.8
| 14 | SUCTION TIP - NO. 15 -Qty.4
| 15 | ADSON FORCEP - NON TOOTHED 12 CM -Qty.4
| 16 | ADSON FORCEP - TOOTHED 12 CM -Qty.4
| 17 | DISSECTING FORCEP - NON TOOTHED 12CM -Qty.4
| 18 | DISSECTING FORCEP - NON TOOTHED 15CM -Qty.4
| 19 | DISSECTING FORCEP - NON TOOTHED 18CM -Qty.4
| 20 | DISSECTING FORCEP - TOOTHED 12CM -Qty.4
| 21 | DISSECTING FORCEP - TOOTHED 15CM -Qty.4
| 22 | DISSECTING FORCEP - TOOTHED 18CM -Qty.4
| 23 | PERIOSTAL ELEVATOR - CURVED -Qty.4
| 24 | PERIOSTAL ELEVATOR - STRAIGHT -Qty.4
| 25 | MOSQUITO ARTERY - CURVED 10 CM -Qty.20
| 26 | MOSQUITO ARTERY - STRAIGHT 10 CM -Qty.20
| 27 | ARTERY FORCEP - CURVED 12 CM -Qty.20
| 28 | ARTERY FORCEP - STRAIGHT 12 CM -Qty.20
| 29 | ARTERY FORCEP - CURVED 15 CM -Qty.10
| 30 | ARTERY FORCEP - STRAIGHT 15 CM -Qty.10
| 31 | ARTERY FORCEP - CURVED 18CM -Qty.10
| 32 | ARTERY FORCEP - STRAIGHT 18CM -Qty.10
| 33 | ALLIS FORCEP - 12 CM -Qty.15
| 34 | ALLIS FORCEP - 15 CM -Qty.15
| 35 | ALLIS FORCEP - 18 CM -Qty.10
| 36 | KOCHERS CLAMP - STRAIGHT 12CM -Qty.10
| 37 | KOCHERS CLAMP - STRAIGHT 15CM -Qty.10
| 38 | KOCHERS CLAMP - STRAIGHT 20CM -Qty.10
| 39 | KOCHERS CLAMP - CURVED 12CM -Qty.10
| 40 | KOCHERS CLAMP - CURVED 15CM -Qty.10
| 41 | KOCHERS CLAMP - CURVED 20CM -Qty.10
| 42 | BABCOCK FORCEP - 15 CM -Qty.4
| 43 | BABCOCK FORCEP - 20 CM -Qty.4
| 44 | NEEDLE HOLDER - HEAVY 18 CM -Qty.4
| 45 | NEEDLE HOLDER - HEAVY 15 CM -Qty.4
| 46 | NEEDLE HOLDER - HEAVY 12 CM -Qty.4
| 47 | NEEDLE HOLDER - FINE TIP 12 CM -Qty.4
| 48 | METZ SCISSOR - CURVED 20 CM -Qty.4
| 49 | METZ SCISSOR - STRAIGHT 20 CM -Qty.4
| 50 | METZ SCISSOR - CURVED 15 CM -Qty.4
| 51 | METZ SCISSOR - STRAIGHT 15 CM -Qty.4
| 52 | MAYO SCISSOR - CURVED 15 CM -Qty.4
| 53 | MAYO SCISSOR - STRAIGHT 15 CM -Qty.4
| 54 | MIXTARD ARTERY - 12 CM -Qty.4
| 55 | MIXTARD ARTERY - 15 CM -Qty.4
| 2 | Operating Environment
| 2.1 | The unit shall be capable of operating in ambient temperature of 15°C-35°C and relative humidity of 10-95%
| 2.2 | The unit shall be capable of being stored continuously in ambient temperature of -5°C -40°C and relative humidity of 10-95%
| 3 | Standards & Safety Requirements
| 3.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 3.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
| 4 | Warranty
| 4.1 | Comprehensive warranty for 2 years.
13 |  | Portable ventilator
| 1 | Description of Function | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1.1 | Unit to be used for adult and child ventilation during cardiopulmonary resuscitation and for ventilation during transportation when necessary.
| 2 | Technical Specifications
| 2.1 | To have frequency control 4 to 100 breaths per minute,
| 2.2 | Tidal volume control 20 - 1500 ml,
| 2.3 | Inflation pressure monitor 0 to 100 cm H2O,
| 2.4 | Air mix control zero to 70% air mixture,
| 2.5 | Adjustable relief pressure with audible alarm 20 to 80 cm H2O
| 2.6 | Add on PEEP facility 0 to 10 or 20 cm H2O.
| 2.7 | To be supplied with a sling to enable the user to carry the unit easily and a patient circuit 1.25m
| 2.8 | long 15mm single bore silicone hose
| 2.9 | Autoclavable
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
Adult and Paediatric reusable, autoclaveable silicon breathing circuits: 02 set each
Reusable Masks ( Medium and Large): 02 set each.
| 3 | Operating Environment
| 3.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 3.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
| 4 | Standards and Safety Requirements
| 4.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
| 4.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
| 5 | User Training
| 5.1 | Must provide user training (including how to use and maintain the equipment).
| 6 | Warranty
| 6.1 | Comprehensive warranty for 2 years after acceptance.
| 7 | Maintenance Service During Warranty Period
| 7.1 | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required.
| 8 | Documentation
| 8.1 | User (Operating) manual in English.
| 8.2 | Service (Technical / Maintenance) manual in English.
| 8.3 | List of important spare parts and accessories with their part numbers and costing.