## RFQ 9-2025
RFQ 09-2025- Specs
1 |  | DEFIBRILLATOR WITH AED MODE
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
To be used in emergency & critical care departments to meets various resuscitation and monitoring needs.
Technical Specifications
Defibrillator with complete accessories operated with AC power supply and internal battery
Defibrillation function:
It shall be a manual defibrillator for external defibrillation
Able to perform synchronized defibrillation and non-invasive pacing therapy.
Defibrillation energy selection:
External biphasic: 50 - 270J
External Paediatric: 2 - 20J
System shall be user friendly, lightweight and easily transportable.
Waveform shape: biphasic.
The defibrillator paddles shall be easily interchangeable among adult, child, infant and internal paddles. It shall come with at least adult and paediatric paddles.
Can be used for neonatal/paediatric and adult defibrillation.
The unit shall be able to perform defibrillation and monitoring by using disposable electrodes.
Shall have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
Recharge time shall not be held longer than 15 seconds before discharge.
Energy charge & discharge and other selection/control buttons shall be available at the paddle handles.
ECG monitoring function:
Shall have a 3-leads ECG, Lead I, II & III, monitoring capability protected from defibrillation by mean of ECG electrodes and through-the-paddles monitoring
With heart rate display and alarms
With Lead-fault indicator
Shall have an integrated thermal printer/ recorder with paper speed of 25mm/sec
General function:
Shall have LCD that displaying at least dual ECG channel, HR, battery status, shock indicator and various data. Bidder to specify size of LCD screen and the no. of waveforms which can be displayed.
Shall have audio and visual alarms. (Please indicate in the next column type of alarms available)
Shall have HR limit and shockable rhythms alarms
Shall have a rechargeable battery when it is fully charged it shall deliver approximately 40 - 50 discharges or 2 hours of continuous ECG monitoring. Bidder to specify the type of battery used and number of discharge and monitoring hour.
Must have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
The unit shall be portable and come with a carrying bag able to keep all required accessories and consumables.
Please indicate the weight in kilogram (KG) of the unit included all accessories and carrying case. It shall be within 8kg
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
· Rechargeable battery, 1 piece on the unit
· Thermal paper x 2 rolls/sets
· Power cord x 1 set
· 3 wire ECG cable x 1 set for ECG monitoring
· Disposable ECG electrodes, 50 pieces
. Paddle for adult and paediatric
· Carry Bag/case x 1 set
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-90%
Thu unit shall be capable of operating in ambient temperature of 20-50 deg C and relative humidity of 80%
Standards and safety requirements
Should be FDA or CE approved product certificate.
User Training
Must provide user operation and service training
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
Certificate of calibration and inspection from factory.
2 |  | Electric ICU Bed Five Movement With Mattress
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
Technical Specifications
1. Bed should have Dual Pedestal Design.
2. Should have built in Bed extender of 200mn for long patients.
3. Tuck away side rails with embedded controls.
4. Electrically operated controls should have lock out facility for electric controls at foot end
5. Backrest & leg section emergency release (CPR)
6. Should have Trendelenbug & covered with bed cover made of AB for the frequent disinfections and washing
7. Synchronized command for bed sprung bas with inclination - backrest – 75 degree and above & Femur rest 25-30 degree
8. Should have electric backrest and pelvic sections with simultaneous movement
9. Should have complying to international standards with acimulated power 1,5 degree 70 VA, electrical current protection class 1, shock protection level type B.
10. Release lever for the side rails with four sectors
11. Protection against water invasion
12. Should have following parameters:
a. Length: > 210 cm
b. Width: > 85 cm
c. Height adjustment: 40-85 cm
d. Back section 0-90
e. Leg Section: 0-25
f. Trendlenberg: 10-20
g. Reverse Trendlenberg: 5-10
13. Should have collapsible /tuck – away side rails
14. Should have double X-ray back rest with Cassette tray
15. Should have facility to fix IV rod at all the four corners and place for fixing accessories
16. Should have castors with central braking system and steering facility
17. Should be radiolucent from Neck though Pelvic region
18. Should have bumpers at all 4 corners
19. Should be provided with:
a. Bed ends fixed to the frame
b. Side rails
c. IV rods
d. Mattress
e. X-ray cassette tray, urine bottle holder, drainage bottle holder
20. Load bearing capacity for the bed should be at least 220kg or more, should have 90 degree upright backrest for chest Imaging.
21. Electrical safety conforms to standards for electrical safety IEC-60601/IS-13450
22. Should have Dual Castors of 18-20 cm
23. Electric shock protection facility
24. Zone controls in Bed to track movement of patient on Bed (Optional)
Operating Environment
The unit shall be capable of operating in ambient temperature of 15°C-35°C and relative humidity of 10-95%
The unit shall be capable of being stored continuously in ambient temperature of -5°C -40°C and relative humidity of 10-95%
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metres in length.
Standards & Safety Requirements
Must submit ISO13485:2003/AC:2007 for Medical Devices AND
CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
Training
Must provide user & service training.
Warranty
Comprehensive warranty for 2 years.
Maintenance Service During Warranty Period
Supplier must ensure planned preventive maintenance (PPM) along with corrective/breakdown maintenance whenever required.
Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
3 |  | Nebulizer Heavy Duty
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
Description of function
Used to administer medication in forms of evipouration to the airways for treating obstructive lung diseases.
Technical specifications
Heavy duty, low noise and compact ultrasonic nebulizer
Durable long life compressor.
To provide pressure up to 2.5 bars.
Should produce particle of size 1-5 micron.
Air delivery rate approx. 15 L/min.
Flow control knob and flow gauge.
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
Air delivery tube 2 nos
Dosage cup 2 nos
Reusable mask for pediatrics 2 nos
Reusable mask for adults 2 nos
Bacterial filter 2 nos
Dust filter
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Environmental factors
The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
The unit shall be capable of operating in ambient temperature of 20-50 deg C and relative humidity of 80%
Standards and safety requirements
Should be FDA or CE approved product certificate.
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English
List of important spare parts, accessories and consumables with their part numbers and costing.
Service (Technical / Maintenance) manual in English
4 |  | Autoclave 20 L
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
| 1 | Description of Function
| 1.1 | Autoclaves are required for sterilizing an object in high temperature and high pressure steam.
| 2 | Operational Requirements
| 2.1 | Microprocessor based electrically heated steam sterilizer
| 3 | System Configuration
| 3.1 | Microprocessor based Autoclave with complete accessories.
| 4 | Technical Specifications
capacity: approx. 20 L.
Stainless steel.
Digital controlled temperature and pressure system.
Steam sterilization, up to 135° C
Digital temperature and pressure gauges.
Safety devices : over heat ( low water cut-off switch , safety valve and release valve)
Exhaust system.
Stainless steel basket.
Double wall case.
| 5 | Accessories, spares and consumables
Spare heating element- 2 set
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Climate ,Temperature , Power supply , Humidity, etc. for Sudan.
Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE or USFDA approved product certificate.
Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials.
| 8 | Warranty
| 8.1 | Comprehensive warranty for 2 years.
| 9 | Maintenance Service During Warranty Period
| 9.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 10 | Installation and Commissioning
| 10.1 | The supplier must accomplish proper commissioning of the item onsite.
| 11 | Documentation
| 11.1 | User and/or service manual shall be supplied in English.
| 11.2 | User (Operating) manual in English.
| 11.3 | Service (Technical / Maintenance) manual in English.
| 11.4 | List of important spare parts and accessories with their part numbers and costing.
| 11.5 | Certificate of calibration and inspection from factory.
5 |  | Autoclave (50 L)
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
| 1 | Description of Function
| 1.1 | Autoclaves are required for sterilizing an object in high temperature and high pressure steam.
| 2 | Operational Requirements
| 2.1 | Microprocessor based electrically heated vertical steam sterilizer
| 3 | System Configuration
| 3.1 | Microprocessor based Autoclave (Vertical Model) with complete accessories.
| 4 | Technical Specifications
capacity: approx. 50 L.
Vertical type.
Stainless steel.
Digital controlled temperature and pressure system.
Steam sterilization, up to 135° C
Digital temperature and pressure gauges.
Safety devices : over heat ( low water cut-off switch , safety valve and release valve)
Automatic controlled sterilization cycle.
2 modes sterilization (121°c - 134°c )
Exhaust system.
Stainless steel basket.
Double wall case.
| 5 | Accessories, spares and consumables
Spare heating element- 2 set
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Climate ,Temperature , Power supply , Humidity, etc. for Sudan.
Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 for Medical Devices AND
| 7.2 | CE or USFDA approved product certificate.
Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials.
| 8 | Warranty
| 8.1 | Comprehensive warranty for 2 years.
| 9 | Maintenance Service During Warranty Period
| 9.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 10 | Installation and Commissioning
| 10.1 | The supplier must accomplish proper commissioning of the item onsite.
| 11 | Documentation
| 11.1 | User and/or service manual shall be supplied in English.
| 11.2 | User (Operating) manual in English
| 11.3 | Service (Technical / Maintenance) manual in English
| 11.4 | List of important spare parts and accessories with their part numbers and costing.
| 11.5 | Certificate of calibration and inspection from factory.
6 |  | Hot Air Oven approx. 40 L
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
| 1 | Description of Function
| 1.1 | Hot Air Oven is required for heating a sample under controlled conditions.
| 2 | Operational Requirements
| 2.1 | Microprocessor based system with PID-temperature controller with integrated auto diagnostic system with fault indicator.
| 3 | System Configuration
| 3.1 | Microprocessor based Hot Air Oven.
| 4 | Technical Specifications
capacity: approx. 40 L.
| 4.1 | External: Stainless Steel Casing: w x h x d: Approx.600 x 600 x 600 mm, insulated stainless steel door with locking and rear zinc-plated steel.
| 4.2 | Interior: w x h x d: Approx. 400mm x 400mm x 400mm. Easy to clean, interior made of stainless steel, with supports on the three sides for three adjustable perforated stainless steel shelves.
| 4.3 | Forced air circulation by quiet air turbine/Fan to ensure uniform temperature.
| 4.4 | Fitted with load indicator and safety thermostat take over indicator lamp. LCD/LED indicator.
| 4.5 | Temperature Variation +/- 1.
| 4.6 | Temperature Range- ambient to 250 oC
| 4.7 | Output available for data acquisition.
| 4.8 | Hot Air Oven shall be mounted on suitable epoxy powder coated support stand having 4 robust 360 deg. swivel lockable castor wheels for easy movement and repositioning.
| 5 | Accessories, spares and consumables
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Power supply: 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs. The power cable must be at least 3 metres long.
| 7 | Standards and Safety Requirements
| 7.1 | Must submit ISO 9001 or ISO 13485:2003/AC: 2007 AND
| 7.2 | CE or USFDA or TUV approved product certificate.
| 8 | User Training
| 8.1 | User training must be provided onsite
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years after acceptance.
| 10 | Maintenance Service During Warranty Period
| 10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
| 11 | Installation and Commissioning
| 11.1 | Supplier must accomplish proper installation and commissioning of the equipment on site.
| 12 | Documentation
| 12.1 | User (Operating) manual in English.
| 12.2 | Service (Technical / Maintenance) manual in English.
| 12.3 | List of important spare parts and accessories with their part number and costing.
| 12.4 | Certificate of calibration and inspection from factory.
7 |  | Obstetric and Gynaecological Set
| No. | Item Specifications | Fill your Specification
Manufacturer Name:
Model No.:
Country of Origin:
| 1 | Technical Specifications
1 Mixing/solution Bowl Qt 3
2 Iodine Cup 6 Oz 3 7/16" X 2"
2 Emesis Basin 12 Oz
2 Instr Tray 12 1/4x7 3/4x2 1/4"
2 Tray Cover For 10-1877
2 Roch Ochsner Forceps Straight 1x2 7 ¼
2 Spencer Wells Forceps Straight 5"
2 Spencer Wells Forceps Straight 8"
2 Rampley Sponge Forceps 9 3/4"
1 Scalpel Handle #4
2 Javerts Placenta Forceps Serrated 91/2
1 Umbilical Scissors 4"
| 2 | Operating Environment
| 2.1 | The system offered shall be designed to store and to operate normally under the conditions include Climate, Temperature, Humidity, etc. for Sudan
| 3 | Standards and Safety Requirements
| 3.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND
| 3.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 4 | User Training
| 4.1 | Not applicable
| 5 | Warranty
| 5.1 | Warranty for 2 years.
| 6 | Maintenance Service During Warranty Period
| 6.1 | Standard warranty conditions are applicable.
8 |  | Examination Couch
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Description of Function
| 1.1 | Examination couch for use of health checkup and treatment of patients.
| 2 | Operational Requirements
| 2.1 | An examination couch with upholstered top in two pieces. Adjustable headrest on gas spring.
| 3 | System Configuration
| 3.1 | Examination couch with mattress.
| 4 | Technical Specifications
| 4.1 | The examination couch shall be made of a solid steel sheet and plate construction with anti-corrosive and antirust treated epoxy powder coating with upholstered top.
| 4.2 | All 4 legs of the bed shall be capped with heavy duty rubber footings.
| 4.3 | Overall size of the table must not be less than 1890mm L x 600mm W x 825mm H
| 4.4 | Strong Mild steel tubular construction epoxy powder coated treated. The top base of machine pressed double bent Mild steel sheet epoxy powder coated treated finish
| 4.5 | Gas spring assisted adjustable backrest of approx. size 450mm L x 310mm H with upholstered top.
| 4.7 | Swinging tray must be attached near headrest for BP apparatus and/or other health checkup minor equipment.
| 4.8 | The mattress shall be foldable and shall be designed to bend with the positioning of the bed when the backrest of the bed is adjusted.
| 4.9 | Bidder shall indicate the weight capacity and the total weight of the mattress in kilogram (kg)
| 4.1 | The mattress shall have mid-firmness, with foam density of approximately 0.55kg/ cubic foot, to avoid that the patient would sink down into foam with antibacterial, antistatic, acid resistance, waterproof and washable vinyl or vinylized nylon cover.
| 4.11 | The joints must be smooth and neat finish.
| 5 | Accessories, spares and consumables
| 5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer, the quantity and details of any items included in this offer which have not been specified in this Technical Specifications Form.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature,Humidity, etc. for Sudan.
| 7 | Warranty
| 7.1 | Comprehensive warranty for 2 years after acceptance.
| 8 | Maintenance Service During Warranty Period
| 8.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
9 |  | Lab Stool
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | Technical Specifications
Revolving metal stool circular seat 38cm Dia.
Constructed with mild steel pipe 25mm Dia chrome coated.
Adjustable height pneumatic gas lift.
Comfortably upholstered with fine leather.
Provides with 2 inc castors.
Height Adjustment - 500mm to 600mm
maximum weight >= 150 kgm
| 2 | Standards and Safety Requirements
| 2.1 | Must submit ISO 9001 or ISO 13485:2003/AC:2007 AND
| 2.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
| 3 | Operating Environment
| 3.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature,Humidity, etc. for Sudan.
| 4 | Warranty
| 4.1 | Comprehensive warranty for 2 years after acceptance.