## Sheet1
RFQ 10-2025- Specs
1 |  | Mobile X-Ray Machine
| No. | Item Specifications | Fill your Specification
| 1 | Technical Specifications
A microprocessor controlled mobile x-ray imaging unit is to include but not limited to the following:
Microprocessor controlled x-ray generator, > 7.5 KW
It should have a digital display of mAs and kV and an electronic timer.
KV range:40kV to 110 kV
mA range: 300 mA or more. Please specify mA and seconds separately and not mAs alone.
Shortest exposure time: I ms.
X-Ray Tube: .
Output should match the output of the generator,
Must have a rotating anode with at least: 2500 rpm and focal spot size should be less than Imm.
Rotating anode tube system, with focal spots , 0.6/1.2 mm of large heat storage capacity, > 90,000 HU
Collimator- Manually adjustable multileaf collimator, rotatable ±90°
The exposure release switch should be detachable with a cord of at least 5 meters
Remote control operating distance > 10 metres
Remote control operating radius- 180 deg
Lightweight manual driven unit, with braking system.
Small source image distance, please specify.
Direct 220-240v/50Hz single phase power line connection with built in line voltage compensation.
Grid( Ratio 6:1) of following sizes should be provided (1) each 12”x15” & 10”x12”
| 2 | User Training:
| 2.1 | The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
| 3 | Warranty
| 3.1 | Comprehensive warranty for 2 years from acceptance.
| 4 | Maintenance Service During Warranty Period
| 4.1 | During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
| 5 | Installation and Commissioning
| 5.1 | The equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
| 6 | Documentation
| 6.1 | User (Operating) manual in English.
| 6.2 | Service (Technical / Maintenance) manual in English.
| 6.3 | List of important spare parts and accessories with their part numbers and costing.
| * | Manufacturer Name:
| ** | Model No.:
| *** | Country of Origin: