## RFQ 6-2025
RFQ 06-2025- Specs
1 |  | C-Arm Machine with UPS
| No. | Item Specifications | Fill your Specification
| I | Manufacturer Name:
| II | Model No.:
| III | Country of Origin:
| 1 | 1 Description of Function
| 1.1 | This equipment is used in orthopaedic fractures for imaging of bone pathology or fractures on display monitor during operation/ reduction of fractures.
| 2 | Operational Requirements For continuous fluoroscopy, image storage and retrieval The system offered shall be a general fluoroscopy/radiology system, it should be a non-digital; non-DICOM compatible type.
| 3 | System Configuration
| 3.1 | X-ray C-Arm Mobile with complete accessories
| 4 | Technical Specifications
| 4.1 | X-Ray Generator
Microprocessor based, high frequency inverter generator
Generator Output: not less than 3kW at 100kV
Fluoroscopic/ Radiographic KV range
Lower limit shall not exceed 40 KV
Higher limit shall not be less than 120 KV
Fluoroscopic mA range
Lower limit shall be ~0.1 mA
Upper limit shall be ~9 mA
| 4.2 | X-Ray Tube
Rotating anode type
Single focal spot, shall not be more than 0.6 mm
Nominal voltage: 110 kV
Anode heat storage capacity not less than 300 KHU
Inherent filtration should be at least 3 mm Al eq
| 4.3 | Collimator
Operator controlled automatic collimation
| 4.4 | C-Arm
· Focus - I.I. Distance shall be at least 100 cm
· Depth shall be ~ 75cm
· Horizontal travel at least 200mm
· Vertical travel at least 450 mm
· Orbital movement shall be ~125°
· Swivel range shall be ~12°
· Rotation about horizontal axis shall be more than +/-180°
| 4.5 | Image Intensifier
· At least 23 cm input screen with direct coupling with camera
· Shall be at least 52 lp/ cm
· Noise reduction, scattered light trap for high contrast dynamics
· CCD camera technology with ABC and AGC control
| 4.6 | TV Monitor
· 2 units LCD monitor side by side for live and reference image
· Shall be at least 43 cm with automatic brightness control
| 4.7 | Image rotation
· Shall be at least 625 scanning lines at 50 Hz
· Trolley for 2 display screens and with the alphanumeric keyboard included
· High resolution and anti glare
| 4.8 | Imaging Modes
· Fluoroscopy mode shall have the following facilities:
· Continuous fluoroscopy with last image hold
· Last image hold with at least two frames image memory
· Continuous fluoroscopy with image acquisition rate: about 20 frame/second.
· Hard disk with image storage capacity of at least 30000 images
· RAM Memory of 256 images
· Mosaic display of 16 images
· Zoom (x 2)
· Measures: at least distances,angles
· Come with CD/DVD/RW drive
| 4.9 | Video printer for B/W thermal printing on paper, size 20 x 25 cm, resolutions about 300 dpi; The video printer can be placed on the monitor trolley
| 4.1 | Indicate here other features and software functions included in this offer
| 5 | System Configuration Accessories, spares and consumables
| 5.1 | Video printer for B/W thermal printing 01 no.
| 5.2 | Sterilizable textile cover and clips, for the X-ray tube and the Cassette holder for 24 x 30 cm
| 5.3 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer.
| 5.4 | Suitable UPS with maintenance free batteries, voltage regulation and spike protection for minimum 30 min. back-up shall be supplied with the system.
| 6 | Operating Environment
| 6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
| 6.2 | Should work on 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs and sockets.
| 7 | Standards and Safety Requirements
The unit offered shall be certified to meeting the relevant requirements of TUV, CE mark (MDD), FDA and/ or any equivalent quality and safety standards.
| 8 | User Training
| 8.1 | On site operational training till the familiarity of the system and satisfaction of end user shall be provided.
| 9 | Warranty
| 9.1 | Comprehensive warranty for 2 years.
| 10 | Maintenance Service During Warranty Period
| 10.1 | Preventive & Corrective Maintenance:
During the warranty period supplier must ensure planned preventive maintenance (PPM) at least 3 nos. in a year along with corrective/breakdown maintenance whenever required.
| 11 | Installation, Inspections and Commissioning
| 11.1 | Supplier must accomplish proper installation and commissioning of the equipment on site.
| 11.2 | Inspections to verify the compliance of the offered equipment as per the specifications
| 12 | Documentation
| 12.1 | User (Operating) manual in English
| 12.2 | Service (Technical / Maintenance) manual in English
| 12.3 | List of important spare parts and accessories with their part numbers and costing.
| 12.4 | Log book with instruction for daily,weekly, monthly and quarterly maintenance checklist.