## Specs - 3-2026
Technical specification - RFQ 3-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Ophthalmology Operating Microscope
1. General Description
The Ophthalmology Operating Microscope shall be a high-precision binocular surgical microscope designed for:
Ophthalmic surgeries including cataract, glaucoma, retinal, and corneal procedures
Microsurgical procedures requiring high magnification and illumination
Adult and pediatric ophthalmic patients
The system shall provide excellent optical quality, ergonomic operation, and integrated illumination to ensure surgical precision and patient safety.
2. Scope of Supply
The microscope system shall include, but not be limited to:
2.1 Optical Components
Binocular head with adjustable interpupillary distance
Magnification changer (continuous or step-wise)
Objective lenses with adjustable working distance
High-resolution optics with distortion-free image
2.2 Illumination System
Integrated LED or Xenon light source
Adjustable illumination intensity
Coaxial illumination for shadow-free field
Fiber optic or built-in illumination cable
2.3 Stand and Mobility
Floor stand or ceiling-mounted configuration
Smooth, multi-directional movement
Lockable brakes or stabilizers for safety
Ergonomic arm for surgeon comfort
2.4 Camera and Documentation (Optional)
Integrated HD/4K camera
Video output for recording and teaching
Footswitch control for focus, zoom, and camera functions
2.5 Accessories
Sterilizable or autoclavable handles
Auxiliary lenses (as applicable)
Footswitch for hands-free operation
Protective dust cover and storage case
3. Technical Specifications
3.1 Optical
Magnification range: ≥4× to 25×
Zoom: Continuous or step-wise
Field of view: Wide enough for microsurgical visualization
Working distance: Adjustable 150–400 mm depending on objective lens
3.2 Illumination
LED/Xenon light source, ≥50,000 lux intensity
Coaxial and optional oblique illumination
Adjustable brightness via control panel or footswitch
3.3 Ergonomics & Operation
Smooth rotation, tilt, and focus
Counterbalanced arm for minimal drift
Hand and foot controls for focus, zoom, and illumination
Compatible with surgical microscope accessories
4. Sterilization & Infection Control
Autoclavable handles and accessories
Smooth, cleanable surfaces
Compatible with hospital infection control protocols
5. Standards & Compliance
The system shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical Safety)
IEC 60601-2-13 (Particular requirements for microscopes)
CE Marking and/or FDA Approval
6. Safety Requirements
Electrical and optical safety per IEC standards
Anti-glare, heat-reducing illumination
Stable stand or ceiling mount with anti-tip protection
Emergency stop or power cutoff for light source
7. Packaging & Labeling
Supplied in protective packaging
Each component labeled for easy installation
Instructions for setup, calibration, and use
8. Maintenance & After-Sales Service
Warranty: minimum 24 months
Availability of spare parts and technical support
Preventive maintenance and calibration services
9. Documentation
Supplier shall provide:
User Manual and Service Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Packing List and Warranty Certificate
10. Training
On-site training for surgeons, nurses, and biomedical staff
Instructions for operation, cleaning, sterilization, and maintenance
2 | A-Scan Ultrasound Biometer (AB Scan) – Ophthalmology
1. General Description
The A-Scan Biometer (AB Scan) shall be a high-precision ophthalmic diagnostic device designed for:
Measuring axial length of the eye
Preoperative assessment for intraocular lens (IOL) calculation
Adult and pediatric patients
The device shall provide accurate, reproducible, and reliable measurements for ophthalmic surgeries, including cataract procedures.
2. Scope of Supply
The device shall include, but not be limited to:
2.1 Main Unit
High-resolution digital A-Scan ultrasound biometer
Handheld or desktop configuration
Integrated probe with single-use or sterilizable covers
Built-in measurement software for axial length, anterior chamber depth, lens thickness
2.2 Accessories
Coupling gel
Disposable or reusable probe covers
Power cable and adapters
Optional: Printer or digital output for report documentation
Protective dust cover
3. Technical Specifications
3.1 Measurement Capabilities
Axial length: 10–40 mm
Resolution: ±0.01 mm
Measurement modes: Contact (applanation) and optional immersion
Automatic averaging of multiple readings
Ability to calculate IOL power using built-in formulas (SRK/T, Holladay, Hoffer Q, etc.)
3.2 Display & Interface
LCD or LED display with high visibility
Intuitive user interface
Digital readout of measurements
Optional integration with PC or EMR system
3.3 Probe Specifications
Frequency: 10 MHz or higher
Ergonomic, lightweight probe
Sterilizable or disposable covers
Stable signal acquisition with minimal noise
3.4 Operational Requirements
Power supply: 220–240V AC, 50/60 Hz or as per hospital standards
Ambient operating temperature: 10–35°C
Portable and compact design for clinic or operating room use
4. Sterilization & Infection Control
Autoclavable probe covers or disposable probe covers
Smooth surfaces for easy cleaning
Compatible with hospital infection control protocols
5. Standards & Compliance
The device shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical Safety)
IEC 60601-2-37 (Ultrasound Equipment)
CE Marking and/or FDA Approval
6. Safety Requirements
Non-toxic, biocompatible materials for probe
Electrical safety per IEC standards
Stable and safe probe handling to avoid ocular injury
Overload and short-circuit protection
7. Packaging & Labeling
Supplied in protective packaging
Clearly labeled components
Instructions for setup, operation, and maintenance
8. Maintenance & After-Sales Service
Warranty: minimum 24 months
Availability of spare parts and technical support
Preventive maintenance and calibration services
9. Documentation
Supplier shall provide:
User Manual and Service Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Packing List and Warranty Certificate
10. Training
On-site training for ophthalmologists, nurses, and biomedical engineers
Instructions for measurement procedures, probe handling, and device maintenance
3 | Phacoemulsification (Phaco) Machine – Ophthalmology
1. General Description
The Phaco Machine shall be a high-precision ophthalmic surgical system designed for:
Cataract surgery (phacoemulsification)
Adult and pediatric ophthalmic patients
Safe, efficient, and precise lens fragmentation and aspiration
The system shall provide adjustable power settings, irrigation/aspiration control, and ergonomic design to ensure surgical safety and efficiency.
2. Scope of Supply
The system shall include, but not be limited to:
2.1 Main Console
Phacoemulsification unit with digital control panel
Adjustable ultrasound energy modes (longitudinal, torsional, or combined)
Integrated irrigation and aspiration system
Footswitch for hands-free control
2.2 Accessories
Sterilizable handpiece(s) compatible with various tips
Phaco tips (straight, angled, and mini-flared)
Irrigation and aspiration tubing sets
Disposable or autoclavable fluidics sets
Protective dust cover and transport packaging
2.3 Optional Accessories
Integrated microscope interface
Digital display for fluidics and energy parameters
Connectivity to EMR or surgical recording systems
3. Technical Specifications
3.1 Ultrasound & Energy Delivery
Modes: Longitudinal, Torsional, or Combined
Adjustable power: 0–100% (or equivalent)
Pulse and burst modes for precise energy delivery
Safety features to minimize corneal endothelial damage
3.2 Irrigation & Aspiration
Adjustable aspiration flow: 0–50 cc/min (or as per manufacturer)
Adjustable vacuum: 0–600 mmHg (or equivalent)
Real-time fluidics monitoring
Occlusion-sensing and automatic vacuum control
3.3 Footswitch & Ergonomics
Multi-function footswitch: phaco, irrigation/aspiration, and coagulation
Ergonomic design for surgeon comfort
Smooth, responsive pedal action
3.4 Display & Interface
Digital touchscreen or panel for parameter control
Real-time monitoring of power, vacuum, and flow
Preset programs for different cataract types
4. Sterilization & Infection Control
Autoclavable handpieces and tips
Disposable or autoclavable fluidics and tubing sets
Smooth surfaces for cleaning and disinfection
Compliance with hospital infection control protocols
5. Standards & Compliance
The system shall comply with:
ISO 13485 (Quality Management System)
IEC 60601-1 (Electrical Safety)
IEC 60601-2-18 (Ophthalmic Surgical Equipment)
CE Marking and/or FDA Approval
6. Safety Requirements
Overload, overheat, and short-circuit protection
Safe energy delivery to minimize ocular trauma
Audible and visual alarms for system errors
Non-toxic, biocompatible materials for patient-contact components
7. Packaging & Labeling
Supplied in protective, shock-resistant packaging
Components labeled for easy identification and use
Instructions for setup, calibration, and operation
8. Maintenance & After-Sales Service
Warranty: minimum 24 months
Availability of spare parts and consumables
Technical support for preventive maintenance, calibration, and repair
9. Documentation
Supplier shall provide:
User Manual and Service Manual
Technical Datasheet / Catalog
Packing List and Warranty Certificate
10. Training
On-site training for ophthalmologists, nurses, and biomedical engineers
Instructions for phaco settings, footswitch use, and fluidics management
Guidance on cleaning, sterilization, and preventive maintenance
4 | Plastic Surgery Set (Reconstructive & Aesthetic Surgical Instruments Set)
1. General Description
The Plastic Surgery Set is designed for use in:
Reconstructive surgery
Aesthetic (cosmetic) procedures
Microsurgical and delicate soft tissue operations
The set shall consist of high-precision, reusable surgical instruments suitable for hospital and operating theatre environments.
2. Scope of Supply
The set shall include, but not be limited to, the following instruments:
Cutting Instruments
Fine Surgical Scissors (Straight & Curved)
Iris Scissors
Metzenbaum Scissors
Grasping & Holding Instruments
Tissue Forceps (with/without teeth)
Adson Forceps
Micro Forceps
Clamping Instruments
Mosquito Hemostats
Small Artery Forceps
Needle Handling
Needle Holders (Fine / Micro)
Retractors
Skin Hooks
Small Retractors (e.g., Senn / Ragnell type or equivalent)
Dissection Instruments
Dissecting Scissors
Elevators
📌 Suppliers may propose additional instruments required for a complete and functional set.
3. Technical Specifications
3.1 Materials
All instruments shall be made of:
High-quality medical-grade stainless steel
Corrosion
Rust
Satin or matte finish to reduce glare
Non-magnetic (preferred for precision work)
3.2 Design & Precision
Fine, delicate tips suitable for microsurgical use
Ergonomic design for prolonged procedures
Smooth, precise alignment of jaws and tips
Anti-slip handles for secure grip
3.3 Performance Requirements
High precision in cutting and tissue handling
Minimal tissue trauma
Smooth operation of joints and locking mechanisms
Long-lasting sharpness for cutting instruments
3.4 Durability
Resistant to frequent use
No misalignment or loosening over time
Maintain performance after repeated sterilization cycles
4. Sterilization & Infection Control
Instruments shall be:
Fully autoclavable (≥ 134°C)
Compatible with low-temperature sterilization (if required)
Resistant to chemical disinfectants
Easy to clean (no hidden cavities)
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Free from toxic or hazardous materials
Smooth edges to prevent accidental injury
Designed to minimize tissue damage
7. Packaging & Labeling
Supplied in:
Sterilization tray or medical-grade case
Each instrument shall be:
Clearly identified (engraved or labeled)
Preserved in a way that prevents damage during transport
8. Maintenance & After-Sales Service
Minimum warranty: 12–24 months
Availability of spare parts
(Re-sharpening)
Technical support availability
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training (If Required)
Basic user training upon request
Instructions for cleaning, handling, and sterilization
5 | Pediatric Surgery Set (Children Surgical Instruments Set)
1. General Description
The Pediatric Surgery Set is designed for surgical procedures in infants and children, including:
General pediatric surgery
Neonatal surgery
Soft tissue and minor surgical procedures
The set shall consist of fine, precise, and size-appropriate instruments to ensure safe handling of delicate pediatric tissues.
2. Scope of Supply
The set shall include, but not be limited to, the following instruments:
Cutting Instruments
Fine Surgical Scissors (Straight & Curved)
Iris Scissors
Delicate Dissecting Scissors
Grasping Instruments
Tissue Forceps (fine, atraumatic, with/without teeth)
Adson Forceps (pediatric size)
Dressing Forceps
Clamping Instruments
Mosquito Hemostatic Forceps (small size)
Pediatric Artery Forceps
Needle Handling
Fine Needle Holders (pediatric size)
Retractors
Small Retractors (e.g., Senn type or equivalent)
Skin Hooks
Additional Instruments
Small Scalpel Handles
Probes and Dilators (if required)
Suction Tips (pediatric size)
📌 Suppliers may include additional items necessary for a complete pediatric surgical set.
3. Technical Specifications
3.1 Materials
Instruments shall be made of:
High-quality medical-grade stainless steel
Resistant to:
Corrosion
Rust
Satin/matte finish to reduce glare
3.2 Design & Size
Specifically designed for pediatric use:
Smaller sizes
Fine tips
Lightweight and easy to handle
Ergonomic and balanced design
Atraumatic edges to minimize tissue injury
3.3 Performance Requirements
High precision in delicate surgical procedures
Smooth and accurate movement of joints
Secure locking mechanisms (where applicable)
Minimal tissue trauma
3.4 Durability
Suitable for repeated use
Maintain alignment and sharpness over time
Resistant to wear and deformation
4. Sterilization & Infection Control
Fully autoclavable (≥ 134°C)
Compatible with:
Low-temperature sterilization
Chemical sterilization methods
Easy to clean and disinfect
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Non-toxic materials
Smooth edges and atraumatic tips
Designed to reduce risk in neonatal and pediatric use
High safety when handling sensitive tissues
7. Packaging & Labeling
Supplied in:
Sterilization tray or medical-grade case
Each instrument shall be:
Clearly identified (engraved or labeled)
Proper protection during shipping
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts
Tool resharpening capability
Technical support availability
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training (If Required)
Basic user training
Instructions for cleaning and sterilization
6 | Vascular Surgery Set (Vascular & Microvascular Instruments Set)
1. General Description
The Vascular Surgery Set is intended for:
Open vascular procedures
Peripheral and central vessel surgery
Microvascular and delicate anastomosis procedures
The set shall consist of high-precision, atraumatic instruments designed to handle delicate blood vessels with minimal tissue damage.
2. Scope of Supply
The set shall include, but not be limited to, the following instruments:
Cutting Instruments
Fine Surgical Scissors (Straight & Curved)
Potts-Smith Scissors (angled, delicate) or equivalent
Micro Scissors
Grasping Instruments
Atraumatic Vascular Forceps
DeBakey Forceps or equivalent
Micro Forceps
Clamping Instruments
Vascular Clamps (various sizes, atraumatic)
Bulldog Clamps
Partial Occlusion Clamps
Needle Handling
Fine Needle Holders
Micro Needle Holders
Retractors
Small and fine retractors
Vessel loops and hooks
Dilators & Probes
Vessel Dilators
Vascular Probes
📌 Suppliers may include additional instruments necessary for a complete vascular surgical set.
3. Technical Specifications
3.1 Materials
All instruments shall be made of:
High-quality medical-grade stainless steel
Resistant to:
Corrosion
Rust
Non-magnetic (preferred for precision work)
Satin or matte finish to reduce glare
3.2 Design & Precision
Atraumatic tips designed for vascular handling
Fine and delicate construction for microvascular work
Smooth jaw alignment with no gaps
Ergonomic handles for precise control
3.3 Performance Requirements
High precision in vessel manipulation and suturing
Minimal trauma to vessel walls
Smooth locking mechanisms for clamps
Consistent clamping force without slippage
3.4 Durability
Maintain precision after repeated use
No misalignment or loosening over time
Resistant to wear and deformation
4. Sterilization & Infection Control
Fully autoclavable (≥ 134°C)
Compatible with:
Low-temperature sterilization
Chemical sterilization methods
Easy to clean and disinfect
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instrument Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic design to protect blood vessels
Free from toxic or hazardous materials
Smooth edges and precision tips
Safe for prolonged and delicate procedures
7. Packaging & Labeling
Supplied in:
Sterilization tray or medical-grade instrument case
Each instrument shall be:
Clearly identified (engraved or labeled)
Protected during transportation
8. Maintenance & After-Sales Service
Warranty: minimum 24 months
Availability of spare parts
The possibility of resharpening precision tools
Technical support availability
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training (If Required)
Basic user training
Instructions for cleaning, handling, and sterilization
7 | Autoclave Sterilizer – 250 Liters (Steam Sterilizer)
1. General Description
The Autoclave Sterilizer (250 Liters capacity) shall be a fully automatic, high-performance steam sterilization system designed for use in:
Hospitals
Operating Theatres
Central Sterile Supply Departments (CSSD)
The unit shall be suitable for sterilizing:
Surgical instruments
Textiles
Medical devices
Laboratory materials
2. Capacity and Chamber
Chamber capacity: Approximately 250 Liters
Chamber configuration: Horizontal
Chamber material:
Stainless Steel AISI 316L or equivalent
Outer jacket:
Stainless Steel AISI 304 or equivalent
Chamber construction:
Designed to withstand high pressure and temperature
Internal surface:
Smooth, polished, corrosion-resistant finish
3. Technical Performance
3.1 Sterilization Parameters
Temperature range:
121°C to 134°C
Pressure:
In accordance with steam sterilization standards
Available sterilization cycles:
Wrapped instruments
Unwrapped instruments
Textiles
Liquids (if applicable)
3.2 Control System
Fully automatic microprocessor-controlled system
Touchscreen display (preferred)
User-friendly interface
Multi-language support (English mandatory)
Pre-programmed cycles and customizable programs
3.3 Monitoring and Recording
Real-time display of:
Temperature
Pressure
Time
Integrated data recording system
Built-in printer or USB/export capability
Storage of sterilization cycle records
4. Safety Features
Overpressure protection system
Over-temperature protection
Safety relief valve
Door safety interlock system (prevents opening under pressure)
Emergency stop button
Audible and visual alarm system
5. Door System
Door type:
Hinged or sliding
Sealing system:
High-quality silicone gasket or equivalent
Locking system:
Automatic or semi-automatic
6. Steam Generation
Built-in steam generator or external steam supply
Automatic water filling system
Compatible with:
Water filter
Distilled or treated water
7. Utilities Requirements
Power supply:
220–240V / 50Hz or 3-phase (as required)
Water supply and drainage connection required
Energy-efficient operation preferred
8. Construction and Installation
Compact and robust design
Suitable for hospital environments
Easy installation and maintenance
Mounted on a stable base or frame
9. Standards and Compliance
The equipment shall comply with:
ISO 13485 (Quality Management System)
ISO 17665 (Moist Heat Sterilization)
IEC 61010 / IEC 60601 (Electrical Safety)
CE Marking and/or FDA Approval
10. Accessories
The system shall include:
Stainless steel loading baskets or trays
Loading trolley (if applicable)
Basic consumables
Spare door gasket
Optional (preferred):
Water treatment unit
11. Maintenance and After-Sales Service
Warranty:
Minimum 24 months
Preventive maintenance program
Availability of spare parts for 10 years
Technical support response time:
Within 24–48 hours
12. Documentation Requirements
The supplier shall provide:
User Manual
Service Manual
Installation and operation requirements
Calibration certificate
CE/FDA certificates
Technical datasheet/catalog
13. Training
On-site training for:
End users (CSSD staff)
Biomedical engineers
Training shall include:
Operation
Cleaning
Basic maintenance
8 | Battery-Operated Cannulated Orthopedic Surgical Drill
1. General Description
The Battery-Operated Cannulated Orthopedic Drill shall be a portable, high-performance cannulated surgical drill intended for:
Orthopedic surgeries requiring guidewire-assisted drilling
(e.g., cannulated screw fixation, hip, knee, and trauma procedures)
Small and large bone drilling procedures using cannulated drill bits
Use in operating theatres and emergency settings
The device shall be ergonomic, lightweight, and compatible with standard cannulated orthopedic systems and accessories.
2. Technical Specifications
2.1 Drill Performance
Speed range: 0–3,500 RPM (variable)
Torque: Adjustable, suitable for cortical and cancellous bone
Forward and reverse rotation
High efficiency with minimal vibration
Precision control for delicate procedures
Cannulated design allowing use over guidewires (e.g., 2.0–4.0 mm K-wires)
High concentricity and minimal runout for accurate over-wire drilling
2.2 Battery System
Rechargeable lithium-ion battery pack
Battery life: ≥ 120 minutes of continuous use
Charging time: ≤ 60 minutes per battery
Indicator for battery status (LED display preferred)
Interchangeable battery system
2.3 Drill Handpiece
Lightweight, ergonomic design
Cannulated (hollow shaft) design compatible with guidewires
Sterilizable parts (autoclavable where applicable)
Compatible with standard cannulated drill bits, reamers, and screw systems
Quick coupling system for drill attachments
Sealed design to prevent fluid ingress into cannulation
3. Accessories (Minimum)
Charger (single or dual battery)
Cannulated drill bit set (various diameters and lengths)
Guidewires (K-wires) – multiple sizes
Drill handle extension/adapter (if applicable)
Carrying/sterilization case
Battery pack(s)
Cannulation cleaning brushes/rods
4. Safety Features
Overload protection
Automatic shut-off when battery is low
Anti-slip ergonomic grip
Forward/Reverse safety switch
Resistant to overheating during prolonged use
Secure locking mechanism for cannulated attachments
5. Sterilization & Infection Control
Autoclavable handpiece covers and drill bits
Cannulated components must be fully cleanable internally
Smooth surfaces for easy cleaning and disinfection
Resistant to hospital-grade disinfectants
6. Standards & Compliance
The device shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical and Safety Requirements)
CE Marking and/or FDA Approval
ASTM F899 / ISO 7740 (Orthopedic Surgical Drills and Accessories)
7. Documentation
Supplier shall provide:
User Manual
Service & Maintenance Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Warranty certificate (minimum 24 months)
8. Training
On-site training for surgeons and biomedical engineers
Operation, guidewire handling, cleaning of cannulation, and maintenance
9 | Orthopedic Surgical Saw (Electric / Battery Operated)
1. General Description
The Orthopedic Surgical Saw shall be a high-performance, precise cutting instrument designed for use in:
Orthopedic surgeries (bone cutting, fracture fixation, joint replacement)
Trauma and reconstructive procedures
Hospital operating theatres and emergency orthopedic settings
The device shall be ergonomic, safe, and compatible with standard saw blades and accessories.
2. Technical Specifications
2.1 Saw Performance
Type: Electric / Cordless battery-operated
Speed range: 0–3,500 RPM (adjustable)
Stroke length: 15–20 mm (or equivalent)
Forward and reverse motion
Smooth operation with minimal vibration
Precise and controlled cutting for cortical and cancellous bone
2.2 Battery / Power System (If Battery Operated)
Rechargeable lithium-ion battery
Continuous operation: ≥ 120 minutes per charge
Charging time: ≤ 60 minutes
Battery indicator (LED display preferred)
Interchangeable battery pack system
2.3 Handpiece
Ergonomic and lightweight design for prolonged use
Sterilizable or autoclavable handpiece covers
Quick coupling system for blades
Smooth, anti-slip grip
3. Blades & Accessories (Minimum)
Set of sterile orthopedic saw blades (various types/sizes)
Blade storage and sterilization case
Charger and spare battery (if battery-operated)
Carrying / sterilization case
Blade adapters / connectors
4. Safety Features
Overload protection
Automatic shut-off on low battery (for cordless saws)
Anti-slip ergonomic handpiece
Blade locking mechanism to prevent accidental detachment
Minimal vibration to reduce user fatigue
5. Sterilization & Infection Control
Autoclavable handpiece covers and blades
Smooth surfaces for easy cleaning and disinfecting
Resistant to hospital-grade disinfectants
6. Standards & Compliance
The device shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical Safety, if electric)
ASTM F899 / ISO 7740 (Orthopedic Surgical Saws and Accessories)
CE Marking and/or FDA Approval
7. Documentation
Supplier shall provide:
User Manual
Service & Maintenance Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Warranty certificate (minimum 24 months)
8. Training
On-site training for surgeons and biomedical engineers
Operation, battery handling, cleaning, and maintenance
10 | Tonsillectomy Surgical Set
1. General Description
The Tonsillectomy Set shall be a complete set of surgical instruments designed for:
Tonsil and adenoid removal
Otorhinolaryngology (ENT) surgical procedures
Both adult and pediatric patients
The instruments shall be precise, atraumatic, and durable, suitable for repeated sterilization and hospital use.
2. Scope of Supply
The set shall include, but not be limited to, the following instruments:
Cutting and Dissecting Instruments
Tonsil Scissors (straight and curved)
Metzenbaum Scissors
Adson or small dissecting scissors
Grasping Instruments
Tonsil or tissue forceps (with/without teeth)
Adson Forceps
Clamping and Hemostatic Instruments
Mosquito forceps
Small artery forceps
Tonsil clamps
Retractors and Holders
Tonsil retractors (e.g., Davis or modified Killian type)
Tongue depressors or tongue blades
Mouth gags (e.g., Boyle-Davis type)
Additional Instruments
Needle holders (fine/pediatric size)
Suction tips (if applicable)
Tongue tie or soft tissue elevator
📌 Suppliers may include additional instruments necessary for a complete functional set.
3. Technical Specifications
3.1 Materials
High-quality, medical-grade stainless steel
Resistant to corrosion and rust
Matte or satin finish to reduce glare
3.2 Design & Precision
Fine tips for atraumatic tissue handling
Ergonomic, anti-slip handles
Smooth joint alignment for scissors and clamps
Suitable for delicate ENT tissue manipulation
3.3 Performance Requirements
Precise cutting and tissue handling
Smooth opening/closing of scissors and clamps
Long-lasting sharpness and durability
3.4 Durability
Suitable for repeated sterilization
No misalignment or loosening of joints
Resistant to wear and deformation
4. Sterilization & Infection Control
Fully autoclavable (≥134°C)
Compatible with chemical disinfectants
Easy to clean with no hidden cavities
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Smooth edges to prevent accidental injury
Non-toxic and biocompatible materials
Atraumatic design to protect delicate oral tissues
7. Packaging & Labeling
Supplied in sterilization tray or medical-grade case
Instruments clearly labeled or engraved
Protected to prevent damage during transport
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts and re-sharpening
Technical support for maintenance
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training
On-site training for surgical staff (optional)
Instructions for handling, cleaning, and sterilization
11 | Esophagoscope Surgical Set
1. General Description
The Esophagoscope Set shall be a complete rigid endoscopic instrument set designed for:
Examination and surgical procedures of the esophagus
ENT and gastroenterology applications
Adult and pediatric patients
The instruments shall be precise, durable, and compatible with standard illumination and suction systems.
2. Scope of Supply
The set shall include, but not be limited to:
Esophagoscopes
Rigid esophagoscope of varying sizes (adult and pediatric)
Blades and straight/curved tubes
Removable obturators
Ancillary Instruments
Biopsy forceps
Foreign body grasping forceps
Esophageal dilators
Suction adapters and tips
Illumination and Visualization
Light source connector compatible with standard operating room light source
Fiber optic illumination cable (autoclavable)
Additional Accessories
Handle and sheath for rigid scope
Storage tray with individual instrument slots
Cleaning brushes and protective covers
📌 Suppliers may include additional instruments necessary for a complete functional set.
3. Technical Specifications
3.1 Materials
High-quality, medical-grade stainless steel
Corrosion and rust-resistant
Smooth, polished surface for patient safety
Fiber optic cables made of durable, sterilizable material
3.2 Design & Precision
Blades and tubes designed for atraumatic esophageal access
Ergonomic handles
Smooth joints and detachable components for easy cleaning
Compatible with suction and illumination systems
3.3 Performance Requirements
Clear visualization of the esophageal lumen
Precise biopsy and foreign body retrieval
Long-lasting structural integrity and illumination
3.4 Durability
Autoclavable and compatible with repeated sterilization
Resistant to mechanical wear and deformation
4. Sterilization & Infection Control
Fully autoclavable instruments and sheaths (≥134°C)
Compatible with hospital-grade chemical disinfectants
Smooth surfaces for efficient cleaning
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Smooth edges and atraumatic design
Non-toxic, biocompatible materials
Compatible with standard OR suction and illumination
7. Packaging & Labeling
Supplied in sterilization tray or medical-grade storage case
Each instrument labeled or engraved for identification
Protective arrangement to prevent damage during transport
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts
Technical support for instrument maintenance
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training
On-site training for surgeons and biomedical staff
Instructions for handling, cleaning, and sterilization
12 | Bronchoscope Surgical Set
1. General Description
The Bronchoscope Set shall be a complete rigid and/or flexible bronchoscopy instrument set designed for:
Examination and therapeutic procedures of the trachea and bronchi
Pulmonology, anesthesia, and ENT applications
Adult and pediatric patients
The instruments shall be precise, durable, sterilizable, and compatible with standard OR light sources and suction systems.
2. Scope of Supply
The set shall include, but not be limited to:
Bronchoscopes
Rigid bronchoscope of varying sizes (adult and pediatric)
Flexible bronchoscope (optional)
Removable obturators and sheaths
Ancillary Instruments
Biopsy forceps
Foreign body grasping forceps
Cytology brushes
Suction adapters and tips
Illumination and Visualization
Light source connector compatible with standard OR or endoscopy light sources
Fiber optic illumination cables (autoclavable)
Camera adapter (optional for flexible bronchoscope)
Additional Accessories
Handle and sheath for rigid scope
Storage tray with individual instrument slots
Cleaning brushes and protective covers
📌 Suppliers may include additional instruments necessary for a complete functional set.
3. Technical Specifications
3.1 Materials
High-quality, medical-grade stainless steel for rigid scopes
Corrosion-resistant, autoclavable components
Smooth, polished surfaces for patient safety
Fiber optic cables made of durable, sterilizable material
3.2 Design & Precision
Atraumatic tips for safe insertion
Ergonomic handles for surgeon comfort
Smooth joints and detachable components for easy cleaning
Compatible with suction, illumination, and optional camera systems
3.3 Performance Requirements
Clear visualization of trachea and bronchi
Accurate biopsy and foreign body retrieval
Long-lasting structural integrity and fiber optic illumination
3.4 Durability
Autoclavable and compatible with repeated sterilization
Resistant to mechanical wear and deformation
4. Sterilization & Infection Control
Fully autoclavable instruments and sheaths (≥134°C)
Compatible with hospital-grade chemical disinfectants
Smooth surfaces for easy cleaning
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic tips and smooth edges
Non-toxic, biocompatible materials
Compatible with standard OR suction and illumination systems
7. Packaging & Labeling
Supplied in sterilization tray or medical-grade storage case
Each instrument clearly labeled or engraved
Protective arrangement to prevent damage during transport
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts
Technical support for instrument maintenance
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training
On-site training for surgeons and biomedical staff
Instructions for handling, cleaning, and sterilization
13 | ENT Tower (Complete Endoscopy System for ENT Surgery)
1. General Description
The ENT Tower shall be a modular, integrated endoscopy system designed for:
Otorhinolaryngology (ENT) surgeries
Diagnostic and operative endoscopic procedures
Adult and pediatric patients
The system shall provide high-definition imaging, illumination, and integrated recording/documentation capabilities, ensuring precision, safety, and efficiency in ENT surgeries.
2. Scope of Supply
The ENT Tower shall include, but not be limited to:
2.1 Imaging & Camera System
HD or Full HD camera head (compatible with rigid and flexible endoscopes)
Image processor and control unit
Light source integration (LED or Xenon)
Video output: HDMI/DVI/VGA for external monitors
2.2 Illumination System
Fiber optic light cables
Adjustable intensity control
Cold light source (LED or Xenon)
Compatible with rigid and flexible ENT scopes
2.3 Display & Documentation
Medical-grade monitor(s) ≥ 24 inch
Recording capability (video and still images)
Optional integration with PACS or hospital network
Touchscreen control panel (preferred)
2.4 Suction & Irrigation
Integrated suction unit
Adjustable flow control
Sterile irrigation system connection
Footswitch control (optional)
2.5 Additional Features
Mobile tower with lockable castors
Multiple drawers and storage for instruments and endoscopes
Cable management system
Emergency backup power (optional UPS integration)
3. Technical Specifications
3.1 Camera & Video
Resolution: HD 1080p or higher
Digital zoom and focus control
Compatible with standard rigid ENT scopes (0°, 30°, 70°)
Compatible with flexible endoscopes (if applicable)
3.2 Light Source
Intensity: Adjustable 10–100%
LED/Xenon with long lifespan (≥5,000 hours for Xenon, ≥50,000 hours for LED)
Quick-connect fiber optic cables
3.3 Tower Construction
Stainless steel or coated aluminum mobile frame
Lockable castors for stability
Shelves for video processor, suction, and storage
4. Sterilization & Infection Control
All patient-contact accessories sterilizable/autoclavable
Smooth, cleanable surfaces
Cable and instrument ports designed for infection control
5. Safety & Reliability
CE/FDA certified components
Overheat protection for light source
Electrical safety per IEC 60601-1
Emergency stop function
Stable tower with anti-tip design
6. Standards & Compliance
The system shall comply with:
ISO 13485 (Quality Management System)
IEC 60601 series (Electrical Safety)
CE Marking and/or FDA Approval
EN 60601-2-18 for endoscopic equipment
7. Documentation
Supplier shall provide:
User Manual and Service Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Packing List
Warranty certificate (minimum 24 months)
8. Training
On-site training for surgeons, nurses, and biomedical engineers
Operation, cleaning, sterilization, and maintenance
14 | Mastoidectomy Surgical Set
1. General Description
The Mastoidectomy Set shall be a complete ENT surgical instrument set specifically designed for:
Mastoidectomy and related ear surgeries
Otorhinolaryngology (ENT) procedures
Adult and pediatric patients
The instruments shall be precise, durable, autoclavable, and suitable for delicate bone and soft tissue manipulation.
2. Scope of Supply
The set shall include, but not be limited to:
Cutting and Dissecting Instruments
Mastoidectomy chisels (various sizes)
Curettes (various sizes, sharp/blunt)
Bone gouges
Elevators (various types)
Grasping and Holding Instruments
Alligator forceps (various sizes)
Tissue forceps (with and without teeth)
Needle holders
Clamps and Hemostatic Instruments
Small artery forceps
Mosquito forceps
Retractors (various sizes, including self-retaining if applicable)
Specialized Instruments
Ear speculum (various sizes)
Suction tips (straight and angled)
Mallets for bone chisels (if applicable)
Micro curettes and picks
Additional Accessories
Sterilization tray with individual instrument slots
Handle adapters for powered instruments (if compatible)
📌 Suppliers may include additional instruments necessary for a complete functional mastoidectomy set.
3. Technical Specifications
3.1 Materials
High-quality, medical-grade stainless steel
Corrosion-resistant and autoclavable
Polished matte or satin finish to reduce glare
3.2 Design & Precision
Fine tips for atraumatic tissue and bone handling
Ergonomic, anti-slip handles
Smooth joint alignment for scissors, forceps, and clamps
Instruments suitable for microscopic or endoscopic ENT surgery
3.3 Performance Requirements
Precise bone and soft tissue manipulation
Long-lasting sharpness and structural integrity
Compatible with powered or manual instrumentation
3.4 Durability
Suitable for repeated sterilization (autoclave ≥134°C)
Resistant to mechanical wear and deformation
4. Sterilization & Infection Control
Fully autoclavable instruments
Compatible with hospital-grade chemical disinfectants
Smooth surfaces for efficient cleaning
5. Standards & Compliance
The set shall comply with:
ISO 13485 (Medical Device Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic tips and smooth edges
Non-toxic, biocompatible materials
Handles designed to prevent slippage
7. Packaging & Labeling
Supplied in sterilization tray or medical-grade storage case
Individual instruments labeled or engraved
Protective arrangement to prevent damage during transport
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts
Technical support for instrument maintenance
9. Documentation Requirements
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Origin
Quality Certificates (ISO / CE / FDA)
Packing List
10. Training
On-site training for surgeons and biomedical staff
Instructions for handling, cleaning, and sterilization
15 | Laser Lithotripsy / Fragmentation System (120 W)
1. General Description
The Laser Lithotripsy System shall be a high-power (≥120 Watts), advanced laser platform designed for urological applications, including:
Lithotripsy of kidney, ureter, and bladder stones
Endoscopic and minimally invasive procedures
Use in adult and pediatric patients
The system shall provide high-efficiency stone fragmentation and dusting, with advanced pulse modulation, reusable fiber compatibility, and comprehensive safety features to ensure optimal clinical outcomes and operator safety.
2. Technical Specifications
2.1 Laser System
Laser Type: Holmium:YAG or equivalent advanced laser technology
Wavelength: Approximately 2,100 nm
Maximum Power Output: ≥ 120 Watts
Energy Range: 0.2 – 6.0 Joules per pulse
Pulse Frequency: Up to ≥ 80–120 Hz (higher preferred)
Pulse Width: Adjustable (short, long, and advanced pulse modes)
Pulse Modulation Technology: For improved stone dusting and reduced retropulsion
Peak Power: Adjustable with stable output performance
Cooling System: High-efficiency cooling system enabling continuous operation
2.2 Fiber Compatibility
The system shall be compatible with standard laser fibers
Reusable laser fibers only (mandatory requirement)
Fiber sizes required: 230 µm, 365 µm, and 600 µm
Quantity: Minimum 3 fibers for each size (total 9 fibers)
Fibers shall be autoclavable and designed for repeated sterilization cycles
Minimum durability: ≥ 20–50 sterilization cycles (manufacturer to specify and guarantee)
Fiber tips shall be re-cleavable/strippable to restore performance after repeated use
Quick-connect fiber interface
Adjustable fiber delivery system
2.3 Operating Modes
Stone fragmentation mode (higher energy, moderate frequency)
Stone dusting mode (lower energy, higher frequency)
Fine dusting mode (very low energy, very high frequency)
Advanced pulse shaping/modulation modes for optimized stone management
3. User Interface & Control
High-resolution touchscreen display
Intuitive and user-friendly interface
Adjustable parameters:
Energy (Joules)
Frequency (Hz)
Pulse width
Real-time display of:
Output power
Pulse frequency
Total delivered energy
Pre-set clinical programs (optional)
Audible and visual alerts
4. Safety Features
Key-lock or user authentication system
Footswitch control for laser activation
Emergency stop button
Laser emission indicators (audible and visual)
Eye protection warning system
Advanced overheat protection with automatic shutdown
System self-diagnostics and fault detection
5. Physical & Operational Requirements
Mobile cart-based system with lockable castors
Compact and ergonomic design suitable for OR use
Power supply: 220–240V AC, 50/60 Hz
Noise level: ≤ 60 dB
Designed for continuous, high-volume clinical use
Operating temperature: 10–35°C
6. Sterilization & Infection Control
Reusable laser fibers must be fully autoclavable
Manufacturer shall provide validated reprocessing instructions including:
Cleaning
Disinfection
Sterilization
Maximum reuse cycles
Fibers shall maintain performance throughout the declared lifecycle
Smooth and cleanable external surfaces
Full compatibility with hospital infection control standards
7. Standards & Compliance
The system shall comply with the following standards:
IEC 60601-1 (Medical Electrical Equipment – General Safety)
IEC 60601-2-22 (Particular requirements for surgical laser equipment)
ISO 13485 (Medical Device Quality Management System)
CE Marking and/or FDA Approval
8. Documentation
The supplier shall provide:
User Manual
Technical Datasheet / Product Catalog
Certificate of Conformity (CE and/or FDA)
Validated reprocessing instructions for reusable fibers
Warranty certificate (minimum 24 months)
9. Training
On-site training for urologists, nurses, and biomedical engineers
Training shall include:
System operation
Laser safety
Fiber handling and maintenance
Cleaning and sterilization procedures
16 | URS (Ureterorenoscope)
Technical Specification
Ureterorenoscope System (Semi-Rigid & Flexible – Adult / Pediatric)
1. General Description
The Ureterorenoscope (URS) System shall be a comprehensive endourology system including semi-rigid and flexible ureteroscopes, designed for:
Diagnostic and therapeutic procedures of the ureter and kidney
Endourological interventions including stone removal, laser lithotripsy, and biopsies
Management of adult and pediatric patients
The system shall provide high-definition visualization, atraumatic access, and excellent maneuverability, with full compatibility with standard endourological accessories.
2. Scope of Supply
The system shall include:
2.1 Semi-Rigid Ureteroscope (Adult)
Outer diameter: 7.5 Fr (Adult size)
Working length: 430–500 mm (or equivalent)
High optical resolution system (rod lens or equivalent)
Straight/rigid shaft for ureteral access
Atraumatic distal tip design
Compatible with laser fibers and endourological accessories
2.2 Flexible Ureteroscope (Adult)
Outer diameter: ≤ 7.5 Fr
Working channel: ≥ 3.6 Fr
Deflection: Up/Down ≥ 270°
High-definition digital or fiber optic imaging system
High flexibility and torque resistance
Atraumatic distal tip
2.3 Pediatric Ureteroscope (Flexible Preferred)
Outer diameter / tip size: 4.5 Fr (Pediatric size)
Ultra-slim flexible design suitable for pediatric anatomy
Working channel: ≥ 3.0 Fr
Deflection: ≥ 270° (if flexible type)
Ultra-atraumatic distal tip for pediatric safety
Optimized for low-pressure irrigation
2.4 Accessories
Sheaths and obturators (adult and pediatric sizes)
Biopsy forceps and stone retrieval baskets (standard & mini sizes)
Laser fiber compatibility:
Adult: 200 – 365 μm
Pediatric: 200 – 272 μm
Guidewires and dilators (adult & pediatric)
Irrigation connectors and flow control systems
Suction adapters
Adapters for camera and light source
2.5 Imaging & Visualization
Compatible with endoscopy camera system
High-definition imaging (HD / Full HD optional)
Fiber optic or digital integrated illumination
High clarity with minimal distortion
Excellent visualization of ureteral and renal anatomy
3. Technical Specifications
3.1 Materials
Medical-grade stainless steel and biocompatible polymers
Corrosion-resistant and durable under repeated sterilization
Smooth atraumatic distal tip
3.2 Design & Precision
Ergonomic handle for precise control
Smooth articulation for flexible scope
High maneuverability in complex anatomy
Compatibility with standard endourological instruments
3.3 Performance Requirements
Clear visualization of ureter and renal collecting system
Precise manipulation through working channel
Smooth insertion and withdrawal with minimal tissue trauma
Stable performance under continuous irrigation
3.4 Durability
Autoclavable (≥134°C) or compatible with low-temperature sterilization
Designed for repeated clinical use
High resistance to mechanical wear
4. Sterilization & Infection Control
Autoclavable components where applicable
Compatible with hospital sterilization protocols
Smooth, cleanable surfaces
Manufacturer shall provide validated reprocessing instructions
5. Standards & Compliance
The system shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical Safety for illuminated systems)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic distal tip for adult and pediatric use
Biocompatible, non-toxic materials
Ergonomic design to reduce operator fatigue
Controlled irrigation compatibility (especially for pediatric use)
7. Packaging & Labeling
Supplied in medical-grade sterilization trays or protective cases
Clear labeling for semi-rigid, flexible, adult, and pediatric scopes
Shock-protected transport packaging
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts for all scope types
Technical support for repair, calibration, and maintenance
9. Documentation
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Conformity (CE / FDA)
Packing List
Warranty Certificate
10. Training
On-site training for urologists, nurses, and biomedical engineers
Training shall include:
Semi-rigid and flexible scope handling
Pediatric endourology techniques
Cleaning and sterilization procedures
Safe use of accessories and laser fibers
17 | Transurethral Resection (TUR) Surgical Set
1. General Description
The TUR Set shall be a complete endourological surgical instrument set designed for:
Transurethral resection of the prostate (TURP) and bladder tumors (TURBT)
Adult urology procedures
Safe, precise, and minimally invasive surgery
The instruments shall provide high precision, durability, and compatibility with standard endoscopic and electrosurgical equipment.
2. Scope of Supply
The set shall include, but not be limited to:
2.1 Resectoscope Components
Resectoscope sheath (22–26 Fr adult size)
Inner sheath with obturator
Telescope (0° or 30°)
Working element (monopolar or bipolar compatible)
Irrigation and inflow/outflow ports
2.2 Ancillary Instruments
Electrode loop (resection loop, cutting loop)
Ball electrode (for coagulation)
Irrigation tubing set
Resectoscope handles and adapters
Obturators and connectors
2.3 Additional Accessories
Light source connection
Camera adapter for video documentation (optional)
Sterilization tray with instrument slots
📌 Suppliers may include additional instruments necessary for a complete functional TUR set.
3. Technical Specifications
3.1 Materials
Medical-grade stainless steel for all metallic components
Durable, corrosion-resistant, autoclavable
Smooth surfaces to prevent tissue trauma
3.2 Design & Performance
Atraumatic obturator tip
Ergonomic handle for precise control
Compatible with monopolar and/or bipolar energy sources
Working channel for smooth electrode passage
Clear visualization of bladder or prostate tissue
3.3 Durability
Autoclavable (≥134°C) or compatible with low-temperature sterilization
Resistant to mechanical wear and repeated use
4. Sterilization & Infection Control
Fully autoclavable components
Smooth, cleanable surfaces
Compatible with hospital infection control protocols
5. Standards & Compliance
The device shall comply with:
ISO 13485 (Medical Device Quality Management System)
IEC 60601-1 (Electrical Safety for connected devices)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic obturator tip to prevent urethral injury
Non-toxic, biocompatible materials
Electrical safety compatible with monopolar/bipolar energy
Ergonomic design to reduce surgeon fatigue
7. Packaging & Labeling
Supplied in sterilization tray or medical-grade case
Individual instruments clearly labeled
Protective arrangement to prevent damage during transport
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts and repair services
Technical support for maintenance
9. Documentation
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Conformity / CE / FDA
Packing List
Warranty certificate
10. Training
On-site training for urologists and biomedical staff
Instructions on handling, cleaning, and sterilization
18 | PCNL (Percutaneous Nephrolithotomy) Forceps
PCNL (Percutaneous Nephrolithotomy) Forceps – Standard & Mini-PCNL Compatible
1. General Description
The PCNL Forceps shall be a specialized endourological instrument set designed for percutaneous nephrolithotomy procedures for kidney stone removal, grasping, extraction, and fragmentation of renal calculi during minimally invasive renal surgery.
The instruments shall be suitable for use in both standard PCNL and Mini-PCNL procedures, and compatible with laser lithotripsy and pneumatic lithotripsy probes.
The system shall be designed for use in adult and pediatric patients.
The instruments shall be durable, highly precise, autoclavable, and suitable for repeated surgical use.
2. Scope of Supply
2.1 Forceps Types
The set shall include, but not be limited to:
Alligator stone retrieval forceps
Grasping forceps (different tip sizes)
Cup / biopsy forceps for renal tissue (optional)
Rigid PCNL forceps for standard access
Mini-PCNL forceps for small access tracts
2.2 Mini-PCNL Compatibility (Mandatory)
Designed for mini percutaneous nephrolithotomy systems
Compatible with small access sheaths (mini-PCNL systems)
Suitable for narrow working channels
Compatible with laser and pneumatic lithotripsy assistance
2.3 Accessories
Ergonomic handle compatible with nephroscopes
Adapter connectors for endoscope/sheath systems
Sterilization tray with individual instrument slots
Compatible with standard PCNL (12–30 Fr) and Mini-PCNL systems (14–20 Fr range)
3. Technical Specifications
3.1 Materials
High-quality medical-grade stainless steel
Corrosion-resistant and autoclavable
Smooth atraumatic tips
3.2 Design & Precision
Ergonomic handle for precise control
Smooth jaw opening and closing mechanism
Strong gripping force without slippage
Compatible with nephroscopes and mini nephroscopes
3.3 Performance Requirements
Strong gripping force without tissue trauma
Smooth maneuverability in renal pelvis and calyces
Reliable operation under irrigation and saline conditions
Compatible with laser and pneumatic lithotripsy environments
Suitable for both standard and mini-PCNL procedures
3.4 Durability
Autoclavable at ≥134°C
Resistant to repeated sterilization cycles
No mechanical loosening over time
4. Sterilization & Infection Control
Fully autoclavable instruments
Compatible with hospital disinfectants
Smooth surfaces for easy cleaning
Manufacturer shall provide validated sterilization instructions
5. Standards & Compliance
The instruments shall comply with:
ISO 13485 (Medical Device Quality Management System)
ISO 7153-1 (Surgical Instruments Materials)
CE Marking and/or FDA Approval
6. Safety Requirements
Atraumatic tips to avoid renal tissue injury
Non-toxic, biocompatible materials
Ergonomic design to reduce operator fatigue
7. Packaging & Labeling
Supplied in medical-grade sterilization tray
Instruments clearly labeled for type and size
Protective packaging to prevent transport damage
Individual slots for each instrument
8. Maintenance & After-Sales Service
Warranty: minimum 12–24 months
Availability of spare parts
Technical support for maintenance and repair services
9. Documentation
Supplier shall provide:
User Manual
Technical Datasheet / Catalog
Certificate of Conformity (CE / FDA)
Packing List
Warranty Certificate
10. Training
On-site training for urologists and biomedical engineers
Training shall include:
Standard PCNL technique
Mini-PCNL technique
Instrument handling and selection
Cleaning, sterilization, and maintenance
Safe use with laser and pneumatic lithotripsy systems
19 | General Surgical Instrument Set
1. General Description
This specification covers a comprehensive general surgical instrument set designed for use in a wide range of surgical procedures, including vascular, cardiovascular, thoracic, abdominal, and soft tissue surgeries.
The instrument set is intended for use in operating rooms and surgical departments to support functions such as tissue dissection, vessel control, atraumatic clamping, surgical exposure, retraction, cutting, and suturing.
The set shall be suitable for repeated use under standard hospital sterilization conditions and shall ensure high precision, safety, and ergonomic handling during surgical procedures.
2. Material of Manufacture
All instruments shall be manufactured from high-quality surgical stainless steel (or equivalent medical-grade material) with the following characteristics:
High corrosion resistance.
Autoclavable and suitable for repeated sterilization cycles.
Non-toxic and biocompatible material.
High mechanical strength and durability
Smooth atraumatic finish to minimize tissue damage
Tungsten Carbide (TC) inserts for needle holders and cutting instruments where applicable
3. Technical Specifications (Instrument Categories & Sizes)
The set shall include instruments in the following categories. All dimensions and configurations are approximate and non-binding, and equivalent standard commercial sizes shall be accepted provided clinical function is maintained.
3.1 Atraumatic Vascular Clamps
Bulldog clamp atraumatic: approx. 70 mm (or equivalent)
Neonatal vascular clamps (angled 30° and 45°): approx. 12 cm (or equivalent)
Atraumatic curved clamps: approx. 21.5 cm and 25.5 cm (or equivalent)
Atraumatic S-shape clamp: approx. 27.5 cm (or equivalent)
Atraumatic angled (Mixter-type) clamp: approx. 28.5 cm (or equivalent)
3.2 Tissue Forceps
Large jaw forceps: approx. 20 cm (or equivalent)
Toothed forceps (10×11): approx. 16 cm (or equivalent)
DeBakey forceps: multiple lengths (16 cm, 20 cm, 24 cm, 30 cm or equivalent range)
3.3 Retractors
Kocher retractor: approx. 40×10 mm, 22 cm (or equivalent)
Czerny double-ended retractor: approx. 17.5 cm (or equivalent)
Cushing hook: approx. 28 cm (or equivalent)
Desmarres vein retractor: blade approx. 16 mm, length 13 cm (or equivalent)
Kelly retractor: approx. 25×38 mm (or equivalent)
3.4 Scissors
Rib scissors (Giertz-Stille): approx. 27 cm (or equivalent)
Micro scissors (straight): approx. 18 cm (or equivalent)
Various surgical scissors (angled, curved, heavy-duty): 17–28 cm range (or equivalent)
DeBakey arteriotomy scissors: approx. 17.5 cm (or equivalent)
3.5 Needle Holders
Micro needle holders: 11 cm with different jaw diameters (or equivalent)
Ryder vascular TC needle holders: 13 cm, 15 cm, 18 cm, 23 cm (or equivalent)
3.6 Vascular Dilators
DeBakey dilators: Ø1.0 mm to Ø5.0 mm, approx. 19 cm (or equivalent graduated set)
3.7 Accessories
Rib spreader (adult): opening approx. 185 mm (or equivalent)
Surgical mallet: approx. 40 mm head, 750 g, 19 cm (or equivalent)
Kidney dishes, bowls, sterilization cases in standard surgical sizes (or equivalent)
4. Quality Standards & Certifications
The supplier shall provide products compliant with the following international standards:
ISO 13485 (Medical Devices Quality Management System)
ISO 9001 (Quality Management System)
CE Mark (European Conformity) or equivalent international certification
FDA approval (if applicable based on origin)
RoHS compliance (where applicable)
Declaration of Conformity (DoC)
5. Warranty
Minimum warranty period: 24 months from date of acceptance
Warranty shall cover manufacturing defects and material failures
Includes replacement or repair of defective instruments
All replacement parts and labor shall be covered during warranty period
Response time for service support: ≤ 72 hours (recommended)
6. Preventive Maintenance & Service
The supplier shall provide:
Guidance for cleaning, sterilization, and maintenance procedures
Availability of spare parts or equivalent instruments for at least 10 years
7. Documentation (Mandatory)
The supplier must submit the following documents:
User, installation, and maintenance manuals
Technical datasheets for all instrument categories
Material composition certificates
Certificate of Origin
CE / ISO certificates
Sterilization and reprocessing guidelines
Declaration of conformity
20 | Inverted Fluorescence Microscope
1. General Description
A reliable inverted fluorescence microscope designed for routine laboratory applications, including:
Cell culture observation (including kidney cells)
Basic immunofluorescence applications
Evaluation of sample uniformity and consistency
Routine imaging and documentation
2. Optical System
Infinity-corrected optical system
Standard diffraction-limited optical resolution
Field of view ≥ 20–22 mm
Inverted configuration (objectives positioned below the stage)
3. Observation Modes
Brightfield
Phase Contrast
Fluorescence (limited to standard channels)
Differential Interference Contrast (DIC) (optional)
4. Objectives
Plan Achromat objectives (Plan Apochromat optional)
Magnifications: 4×, 10×, 20×, 40× (60× optional)
Numerical aperture suitable for routine imaging (up to approximately 1.3)
5. Fluorescence System
LED illumination source with long lifetime (≥ 30,000–50,000 hours)
Support for standard fluorescence channels (e.g., DAPI, FITC, TRITC)
Filter cubes with appropriate dichroic mirrors
Manual filter selection (motorized optional)
6. Stage and Focus System
Mechanical XY stage
Compatible with culture dishes and multiwell plates
Coarse and fine focusing mechanisms
Stable and precise manual focusing system suitable for routine use
7. Imaging System
Digital camera (CMOS or CCD)
Resolution ≥ 5 megapixels
Camera with high sensitivity and low noise suitable for fluorescence imaging
Global shutter or rolling shutter with fast readout capability suitable for live imaging applications
Minimum exposure time ≤ 1 ms
Live image display capability
Image capture and storage functionality
Basic image acquisition features
8. Environmental Conditions
Operation under standard laboratory conditions
No live-cell environmental control required
9. Software
Image acquisition software
Multi-channel fluorescence imaging support
Basic image analysis (measurement, annotation, optional cell counting)
Compatible with standard operating systems
10. Electrical Requirements
Power supply: 220–240 V AC, 50/60 Hz
11. Safety Requirements
UV protection shielding
Automatic light shut-off feature
Compliance with IEC 61010 or equivalent safety standards
12. Quality and Compliance
ISO 9001 certification
CE marking
RoHS compliance
ISO 13485 and FDA (if applicable)
13. Documentation (Mandatory)
The supplier shall provide:
User manual
Installation manual
Basic service documentation
Technical datasheet
Calibration information
Copies of relevant certificates
14. Warranty (Mandatory)
Minimum warranty period: 24 months from the date of installation
Warranty shall include:
All parts and labor
On-site service and support
Replacement of defective components
Response time ≤ 72 hours
15. Training (Mandatory)
On-site training (1–2 days) covering:
End users (operation and imaging)
Technical staff (basic maintenance)
Training materials must be provided
16. Preventive Maintenance and Service
Basic preventive maintenance schedule
Availability of spare parts for at least 5 years
Maintenance records recommended
Optional Annual Maintenance Contract (AMC)
17. Accessories
Computer workstation
Calibration slide
Standard fluorescence filter sets
21 | Hemodialysis Water Treatment Plant (HD-WTP)
Tender Document for 10 Dialysis Machines
1. General Description
The Contractor/Supplier shall design, manufacture, supply, install, test, commission, validate, and hand over a complete Hemodialysis Water Treatment Plant (HD-WTP) intended to produce ultrapure water suitable for hemodialysis applications.
The system shall be designed to serve a dialysis unit consisting of ten (10) hemodialysis machines with continuous, reliable, and safe operation in accordance with international dialysis water quality standards.
The supplied system shall include all equipment, accessories, piping, instrumentation, controls, electrical components, cooling systems, disinfection systems, and all necessary items required for complete and fully functional operation.
The HD-WTP shall be suitable for:
Conventional Hemodialysis (HD)
High-Flux Dialysis
Hemodiafiltration (HDF), where applicable
The complete system shall be designed for:
Continuous 24/7 recirculation
High microbiological safety
Ease of maintenance
Energy-efficient operation
Long-term operational reliability
2. Applicable Standards
The system shall comply with the latest editions of the following standards and guidelines:
Association for the Advancement of Medical Instrumentation Standards:
ANSI/AAMI RD52
ANSI/AAMI RD62
ANSI/AAMI/ISO 23500 Series
International Organization for Standardization:
ISO 23500 Parts 1–5
European Pharmacopoeia Requirements
World Health Organization Drinking Water Guidelines
CE MDR or equivalent international medical device certification
ISO 13485 certification for medical device manufacturing
3. System Capacity & Design Basis
The system shall be designed based on the following minimum criteria:
Number of dialysis machines :10 units
Average water consumption per machine :500–800 L/day
Estimated total daily demand :5–8 m³/day
Design safety margin :Minimum 30%
Total design capacity: 7–10 m³/day
RO production capacity : Minimum 500–700 L/hour
Operating duration : Minimum 16–20 hours/day
System operation : Continuous recirculation (24/7)
The Supplier shall verify all design calculations based on actual site conditions and raw water analysis.
4. Feed Water Requirements
The Supplier shall review available raw water analysis prior to final design.
The following parameters shall be considered during system design:
Hardness
Iron
Manganese
TDS
Turbidity
Free chlorine/chloramines
Silica
SDI
Microbiological contamination
If no water analysis is available, the Supplier shall clearly state all assumed feed water conditions.
5. Final Water Quality Requirements
The HD-WTP shall continuously achieve dialysis water quality compliant with AAMI/ISO standards.
Product Water Quality
Bacteria :< 0.1 CFU/mL
Endotoxins: < 0.03 EU/mL
Free chlorine: 0 mg/L
TOC :< 0.5 mg/L
Product water conductivity :≤ 5 µS/cm preferred
Alarm conductivity: >10 µS/cm
Online Monitoring (Mandatory)
Continuous online monitoring shall include:
Conductivity
Pressure
Flow rate
Temperature
Free chlorine
Tank level
Loop return conductivity
Monitoring locations shall include:
Post-RO Pass 1
Post-RO Pass 2
Storage tank
Distribution loop return
6. Pretreatment System
The pretreatment section shall include all necessary stages required to protect the RO membranes and ensure reliable operation.
6.1 Multimedia Sand Filter
The multimedia filter shall provide high-efficiency suspended solids and turbidity removal.
Requirements
Automatic backwash operation
Differential pressure monitoring
FRP or SS316L pressure vessel
Automatic control valve
Sampling ports
6.2 Iron Removal Filtration System (Mandatory)
A dedicated iron removal system shall be provided as a mandatory pretreatment stage.
Performance Requirements
Outlet iron concentration:<0.05 mg/L preferred
Maximum allowable :≤0.1 mg/L
Media
Acceptable media include:
Birm
Manganese Greensand
Katalox Light
Equivalent catalytic media
Configuration
FRP or SS316L pressure vessel
Fully automatic operation
Automatic backwash
PLC integration
Sampling ports (inlet/outlet)
Installation Location
The iron removal filter shall be installed:
After multimedia filtration
Before activated carbon filtration
6.3 Activated Carbon Filtration System (Safety Critical Stage)
A catalytic activated carbon filtration system shall be provided for complete removal of chlorine and chloramines.
Configuration
Preferred:
Dual Vessel (Lead/Lag)
Acceptable:
Single Vessel (subject to approval)
Requirements
Catalytic activated carbon media
EBCT ≥10 minutes at peak flow
Automatic backwash
Proper internal distribution
Mandatory Safety Features
Online free chlorine analyzer
Continuous monitoring
Audible and visual alarms
PLC integration
Performance Requirement
Outlet free chlorine concentration:
0 mg/L at all times
6.4 Water Softener
The softener system shall include:
Duplex configuration
Automatic regeneration
Continuous operation
Online hardness monitoring
Performance
Residual hardness:
<1 ppm as CaCO₃
RO shutdown shall occur upon hardness breakthrough.
6.5 Chemical Dosing System (Optional)
Where required, anti-scalant dosing shall be provided.
The dosing system shall:
Be interlocked with RO operation
Include chemical dosing pump
Include dosing tank and level monitoring
7. Reverse Osmosis (RO) System
Configuration
Dual Pass RO System (Mandatory)
Capacity
Minimum production:
500–700 L/hour
General Requirements
Membrane rejection ≥99%
SS316L high-pressure pumps
Duty/standby pump configuration
Automatic flushing system
CIP system
Online conductivity monitoring for each pass
Membrane integrity testing
Performance
Recovery rate :60–75%
SDI at RO inlet:<3
8. Post-Treatment System
The post-treatment system shall include:
UV sterilization (254 nm)
Endotoxin-retentive ultrafiltration (0.01 micron)
Optional ozone system (offline use only)
The ultrafilter shall be installed:
After UV unit
Before distribution loop
9. Storage & Distribution System
9.1 Storage Tank
The storage tank shall be sanitary SS316L construction.
Requirements
Capacity: 3000–5000 liters
Fully drainable conical bottom
0.2 µm hydrophobic vent filter
Spray ball for CIP
Level transmitter
Overflow protection
9.2 Distribution Loop
General Requirements
Continuous recirculation (24/7)
Ring loop configuration
SS316L electropolished piping
Hygienic sanitary fittings
No dead legs permitted
Design Criteria
Velocity:1.0–1.5 m/s
Loop turnover:6–10 times/hour
Compatibility
The loop shall be compatible with:
Thermal disinfection (80–85°C)
Chemical disinfection
10. Chiller System (Mandatory)
A dedicated industrial chiller system shall be provided for cooling the dialysis water distribution loop and storage system to minimize microbiological growth.
Performance Requirements
Loop temperature :20–25°C adjustable
Maximum allowable temperature :≤25°C
Ambient design condition :45–50°C
Chiller Components
The system shall include:
Air-cooled chiller
SS316L plate heat exchanger
Circulation pump
Expansion tank
Digital temperature controller
Monitoring & Control
The system shall include:
Online temperature monitoring
High-temperature alarm
PLC integration
Automatic operation
Thermal Disinfection Compatibility
The chiller shall automatically isolate or bypass during thermal disinfection cycles.
11. Disinfection System
The system shall include automated disinfection capability for:
RO system
Storage tank
Distribution loop
Methods
Preferred:
Thermal disinfection
Chemical disinfection:
Peracetic acid
Citric acid
Features
Automated cycles
Validation and logging
Safety interlocks
12. Control & Automation System
The system shall include a PLC-based control system with HMI touchscreen interface.
Monitoring Parameters
Pressure
Flow
Conductivity
Temperature
Free chlorine
Chiller status
Alarm status
Features
Alarm and fault logging
Data logging (minimum 12 months)
Temperature trending
Optional remote monitoring
User access control
13. Electrical System
Requirements
Power supply:
220–240V / 380–415V
50–60 Hz
MCC panel
Surge protection
UPS backup:
Minimum 60 minutes
14. Materials of Construction
Wetted Parts
Mandatory materials:
SS316L stainless steel
PVDF
Medical-grade PVC
Gaskets
EPDM
PTFE
Not Permitted
Galvanized steel
Industrial PVC
Low-grade plastics
Carbon steel in wetted areas
15. Compliance & Validation
The Supplier shall provide:
IQ / OQ / PQ validation
FAT
SAT
Water quality validation reports
Temperature validation
Continuous operation trial
Compliance Standards
ISO 13485
CE MDR (or equivalent)
AAMI / ISO 23500
Pharmacopoeia standards
16. Documentation Requirements
The Supplier shall provide:
PFD
P&ID
General Arrangement Drawings
Electrical Schematics
O&M Manuals
Spare Parts List
Validation Reports
PLC/SCADA Backup
Chiller calculations
As-built drawings
17. Warranty
Minimum Warranty Period
24 months from commissioning
Coverage
Mechanical components
Electrical components
Pumps
Instrumentation
PLC/control system
Conditions
Repair or replacement
Response time: 48–72 hours
18. Training Requirements
The Supplier shall provide on-site training covering:
Operation
Monitoring
Maintenance
Disinfection
Troubleshooting
Training Duration
Minimum 3–5 days
19. After-Sales Service
The Supplier shall provide:
Technical support
Emergency support
Spare parts availability for minimum 10 years
Optional Annual Maintenance Contract (AMC)
20. Vendor Qualification Requirements
The bidder shall demonstrate:
Experience with similar dialysis water systems
Availability of local technical support
Availability of spare parts and consumables
Qualified service personnel
22 | Technical Specification HIGH PERFORMANCE LIQUID CHROMATOGRAPHY (HPLC) SYSTEM
HIGH PERFORMANCE LIQUID CHROMATOGRAPHY (HPLC) SYSTEM
Technical Specifications – Tender Document (Final Formatted Version)
1. General Requirements
The supplier shall provide a brand-new, fully integrated HPLC system (not refurbished or used).
The system shall be suitable for pharmaceutical quality control (QC), chemical analysis, and laboratory applications.
The system shall include all necessary modules for complete operation including pump, autosampler, detector, column oven, degasser, software, PC, and accessories.
2. Pump System
Quaternary pump system or equivalent capable of handling at least four mobile phases (A, B, C, D)
Flow rate range: 0.001 to 10.0 mL/min or better
Maximum operating pressure: 5000 psi or higher
Flow accuracy: ±0.1% or better
Gradient formation with automatic online solvent blending technology
Suitable for isocratic and gradient elution modes
Automatic solvent conditioning including priming, purging, and equilibration functions
3. Autosampler
Fully automated sample injection system
Sample capacity: minimum 120 vials or more
Injection volume: 0.1 to 100 µL programmable
Injection precision: RSD ≤ 0.5%
Sample carryover: not more than 0.0025% (using caffeine test compound) under validated conditions
Sample compartment temperature control: 4°C to 40°C or better
Random access and continuous unattended operation capability
4. UV/Vis Detector
UV/Vis variable wavelength detector (VWD) or equivalent
Wavelength range: 190 to 1000 nm or better
Wavelength accuracy: ±1 nm or better
Wavelength repeatability: ±0.1 nm or better
Wavelength resolution (readability): 0.1 nm or better
Spectral bandwidth: ≤ 5 nm
Baseline noise: ≤ 4 × 10⁻⁶ AU or better
Baseline drift: low drift for stable operation
Photometric measuring range: 0.0001 to 4.0000 AU or better
Linearity: 0 to 2.5 AU or better
Low-volume flow cell suitable for high sensitivity and pressure resistance
Fast response time with automatic baseline correction
5. Column Oven
Temperature range: ambient +5°C to 80°C or better
Temperature stability: ±0.5°C or better
Suitable for single or multiple column operation
Digital temperature control system
6. Degasser
Online vacuum degasser system
Capable of degassing at least 4 solvent channels
Continuous removal of dissolved gases from mobile phases
Improves baseline stability and reduces detector noise
7. Column Capacity
Supports at least one standard analytical column
Expandable configuration supporting up to three columns
Compatible with standard HPLC columns (e.g., 4.6 mm ID columns)
Optional column switching capability
8. Control System & Software
External PC-based control system
Fully integrated Chromatography Data System (CDS) software
Software such as Waters Empower or equivalent
Full instrument control (pump, detector, autosampler, oven)
Data acquisition, processing, and reporting functions
Method development and storage capabilities
Audit trail and user access control compliant with GMP/GLP requirements
9. Performance Requirements
Retention time repeatability: RSD ≤ 1%
Peak area repeatability: RSD ≤ 2%
Stable baseline during continuous operation
System suitability compliance for pharmaceutical QC applications
10. Installation, Qualification & Validation
Full installation and commissioning at site by the supplier
Qualification shall include:
Installation Qualification (IQ)
Operational Qualification (OQ)
Performance Qualification (PQ)
Full documentation and validation reports shall be provided
11. Electrical & Environmental Requirements
Power supply: 220–240 V, 50/60 Hz
Operating temperature: 10°C to 35°C
Relative humidity: up to 85% non-condensing
Suitable for continuous laboratory operation under standard conditions
12. Electrical Safety Certification
The system shall comply with IEC 61010-1 or equivalent international safety standards for laboratory equipment.
The instrument shall be supplied with a valid electrical safety compliance certificate issued by the manufacturer or an accredited certification body.
The system shall ensure protection against:
Electric shock hazards
Over-voltage and over-temperature conditions
Internal electrical faults
The equipment shall be safe for continuous operation in a laboratory environment.
13. Quality Certifications
The system shall be brand new (not refurbished)
Manufacturer certified under ISO 9001 or equivalent quality management system
Calibration certificates for major components shall be provided
Compliance with international HPLC manufacturing standards
Certificates of origin and conformity shall be included
14. Training
On-site training for 3–5 users (analysts/technicians)
Training shall include:
System operation
Maintenance procedures
Software and data analysis
Method development and troubleshooting
Training duration: 3–5 working days minimum
Training manuals and materials shall be provided
15. Warranty
Minimum warranty period of 24 months (2 years)
Covers all major components including pump, detector, autosampler, and electronics
Free repair or replacement of defective parts due to manufacturing faults
Technical support response time not exceeding 72 hours
16. Maintenance & After-Sales Support
Spare parts availability for at least 10 years
Preventive maintenance services included or available
Software updates and technical support provided
Authorized local or regional service support required
Diagnostic tools and service manuals included
17. Required Documents
The supplier must submit the following documents:
Technical data sheets (TDS)
Manufacturer catalogues and brochures
Certificate of origin
ISO quality certificates
Compliance statement with specifications
Operation and maintenance manuals
Software documentation
Training plan
Spare parts list
After-sales service plan
Warranty certificate from manufacturer or authorized distributor
23 | Flexible & Rigid Ureteroscope System (Adult/Pediatric)
1.1 Flexible Ureteroscope (Adult & Pediatric)
Outer diameter: ≤ 7.5 Fr
Working channel: ≥ 3.6 Fr
Deflection: ≥ 270° up/down
Digital HD / 4K imaging system
Laser compatible (200–365 µm)
1.2 Semi-Rigid Ureteroscope (Adult)
Outer diameter: 6.5 Fr
Working length: 430–500 mm
Rod lens optical system
Laser compatible
1.3 Pediatric Flexible Ureteroscope
Outer diameter: 4.5–6 Fr
Working channel: ≥ 3.0 Fr
Deflection: ≥ 270°
Low-pressure irrigation compatible
1.4 Pediatric Ureteroscope (Ultra-Slim)
Outer diameter: 4.8 – 6 Fr
HD / 4K digital imaging system
Laser compatible
2. CYSTOSCOPY SYSTEM
Rigid cystoscope set (2 sizes 19 & 21 fr )
Telescope compatibility: 0°
Continuous flow irrigation
Complete sheath, obturator, bridges, and adapters
3. RESECTOSCOPY SYSTEM
3.1 Monopolar Resectoscope Set (26 Fr)
Working element: 26 Fr
Sheath: 24 Fr
Telescope compatibility: 30°
Monopolar system
Visible obturator
Continuous flow irrigation
3.2 Pediatric Monopolar Resectoscope (11 Fr)
Cutting loops & hook electrodes
HF monopolar cable included
3.3 Monopolar Cutting Loops
Standard, round, angled types
Compatible with resectoscope system
4. URETHROTOMY SYSTEM
4.1 Urethrotome Set (21 Fr)
Optical internal urethrotome
Cold knife / electrosurgical compatible
Complete set
5. NEPHROLOGY / PCNL SYSTEM
5.1 Nephroscope (22 Fr)
Rigid nephroscope
Wide working channel
Laser / ultrasonic compatible
5.2 Complete PCNL System
Includes:
Access sheaths (16–30 Fr)
Dilators
Guidewires
Stone forceps
Irrigation system
Lithotripsy compatibility
6. ACCESSORIES
6.1 Grasper Forceps (5 Fr)
Stone retrieval forceps
Flexible shaft
6.2 Grasper Forceps (3 Fr, for URS)
Ultra-fine design
Compatible with flexible URS
6.3 Punch Cup Biopsy Forceps
For bladder and urethral biopsy
6.4 Stone Punch Set (24 Fr)
Murmayer type stone crushing system
7. TELESCOPES
7.1 Telescope 30° (4 mm, 302 mm) – Qty 2
7.2 Telescope 0° (4 mm, 302 mm) – Qty 2
7.3 Pediatric Telescope 30° compatible with the above set
7.4 Pediatric Telescope 0° compatible with the above set
Rod lens system
HD / 4K compatible
Fully autoclavable
8. LASER & ENERGY / DIATHERMY SYSTEMS
8.1 Laser unit
Holmium Laser above 35 WATT ,for stone .
8.3 Monopolar HF Cable
Fully insulated
Electrosurgical compatible
8.4 Ligasure / Diathermy System
Bipolar vessel sealing system for endoscopic and open urological procedures
Adjustable power settings for tissue sealing and cutting
Atraumatic, ergonomic handpiece
Footswitch and connection cables included
Compatible with sterilization protocols (autoclavable components where applicable)
Designed for hemostasis, tissue dissection, and precise coagulation
Manufacturer instructions included for safe use and reprocessing
Future-ready to integrate new tools or software updates
9. Low-pressure irrigation system
10. IMAGING SYSTEM
HD / Full HD / 4K compatibility
Digital or fiber optic system
High illumination intensity
Minimal image distortion
Accurate color reproduction
Fully compatible with all scopes and accessories
Sterilizable components where applicable
Future-proof for AI-assisted imaging and future upgrades