## Specs - 5-2026
Technical Datasheet RFQ - 05/2026
| Technical Specifications | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Hemodialysis Machine (Dialysis Unit)
General Requirements
1. The unit shall be a fully automatic hemodialysis machine designed for the treatment of acute and chronic renal failure in hospital and dialysis center settings.
2. The machine shall comply with international medical device standards, including IEC 60601-1 and relevant dialysis-specific standards.
3. The system shall support single-needle and double-needle dialysis modes.
4. The machine shall be microprocessor-controlled with real-time monitoring of all treatment parameters.
5. The unit shall be compatible with a central Reverse Osmosis (RO) water treatment system.
Technical Specifications
1. Dialysis Modes: Bicarbonate dialysis, acetate dialysis, and isolated ultrafiltration (UF) mode.
2. Blood Flow Rate (Qb): 50 – 600 ml/min adjustable.
3. Dialysate Flow Rate (Qd): 300 – 800 ml/min adjustable.
4. Ultrafiltration Control: Volumetric controlled system with programmable UF rate.
5. UF Accuracy: ±1% or better.
6. Dialysate Temperature: 35°C – 39°C adjustable.
7. Conductivity Monitoring: Continuous automatic monitoring with alarm system.
8. Heparin Pump: Integrated programmable syringe pump.
9. Blood Pump: Precision roller pump.
10. Air Bubble Detector: Ultrasonic detection system with automatic venous line clamp.
11. Blood Leak Detector: Optical sensor system.
12. Pressure Monitoring: Continuous arterial, venous, and TMP monitoring.
13. Urea Clearance / Dialysis Adequacy: The system shall allow user-adjustable treatment parameters to optimize dialysis dose (e.g., Kt/V target or equivalent adequacy control settings).
Bicarbonate Delivery System
1. The machine shall be equipped with a dry bicarbonate delivery system in powder form (non-liquid concentrate system).
2. The system shall utilize a closed cartridge or dry powder container system for on-site preparation of bicarbonate dialysate.
3. Pre-mixed liquid bicarbonate concentrate systems shall NOT be accepted.
4. The system shall ensure automated mixing and preparation under closed conditions.
5. The system shall be fully compatible with the dialysis machine and approved by the manufacturer.
Control and Display System
1. Large LCD or touch screen display (minimum 10 inches).
2. Real-time display of all treatment parameters.
3. Multilingual user interface (including English).
4. Full alarm system for all critical parameters.
5. Event logging and treatment data storage capability.
Safety Features
1. Automatic shutdown in case of critical alarms.
2. Venous air protection system with automatic clamp.
3. Built-in self-test before each session.
4. Battery backup for short-term power failure.
5. Full compliance with infection control and electrical safety standards.
Water & Dialysate System
1. Compatible with central RO water system.
2. Integrated dialysate proportioning system.
3. Automatic rinse, cleaning, and disinfection cycles.
4. Compatible with standard disposable bloodlines and dialyzers.
Accessories
1. Blood pressure cuff.
2. RO connection sets.
Electrical Requirements
1. Power supply: 220–240V, 50/60 Hz.
2. Energy-efficient operation mode.
Certifications & Standards
1. CE Mark (EU MDR 2017/745).
2. ISO 13485.
3. IEC 60601-1.
4. IEC 60601-2-16.
5. ISO 14971.
6. ISO 10993.
7. FDA approval (if applicable).
8. Certificate of Origin and Factory Acceptance Test (FAT) reports.
Documentation (Required Documents)
The supplier shall submit the following documents with the bid:
1. Original product brochure/catalogue with full technical specifications.
2. Manufacturer’s compliance declaration to all required standards.
3. CE Certificate (EU MDR 2017/745) or equivalent international certification.
4. ISO 13485 Quality Management System Certificate.
5. IEC 60601-1 and IEC 60601-2-16 compliance certificates.
6. ISO 14971 Risk Management Certificate.
7. ISO 10993 Biocompatibility Certificate (blood-contact materials).
8. Certificate of Origin issued by the manufacturer or authorized chamber of commerce.
9. Factory Acceptance Test (FAT) report.
10. Installation, commissioning, and acceptance test protocol.
11. User and service manuals in English (hard copy and digital).
12. Warranty certificate clearly stating duration and coverage.
13. Maintenance and service agreement proposal (preventive and corrective).
14. Training program details for medical, nursing, and biomedical staff.
15. List of recommended spare parts for at least 5 years of operation.
Maintenance Requirements
Preventive Maintenance (PM)
Every 3 months minimum.
Calibration, inspection, cleaning, and performance testing.
Detailed maintenance reports required.
Corrective Maintenance (CM)
Response within 24–48 hours.
Repair and replacement using original spare parts only.
Comprehensive Maintenance Contract (Optional)
Includes PM, CM, spare parts, and software updates.
Spare Parts & Support
Availability for 10 years.
Local and remote technical support.
Warranty
Minimum 2 years warranty.
Covers preventive and corrective maintenance.
Ensures availability of spare parts and technical support
2 | Hemodialysis Chair (Dialysis Patient Chair)
Technical Specifications
General Requirements
1. The chair shall be a fully motorized medical hemodialysis chair designed for use in dialysis centers and hospital renal units.
2. It shall be suitable for long-duration treatments (minimum 4–6 hours continuous use or more).
3. The frame shall be constructed from high-strength steel with anti-corrosion coating.
4. Upholstery shall be medical-grade, seamless, antibacterial, waterproof, and easy to clean and disinfect.
5. The chair shall provide high stability, patient comfort, and safe positioning throughout the dialysis session.
Technical Specifications
1. Type: Fully motorized electric dialysis chair.
2. Motors: Minimum 3–5 electric motors for independent adjustment of:
Backrest
Leg rest
Height
Trendelenburg and Reverse Trendelenburg positions
3. Load Capacity: Not less than 180–250 kg.
4. Full Flat Position: The chair shall allow a fully horizontal position suitable for CPR and emergency procedures.
5. Seat Width: Minimum 55–70 cm.
6. Total Length: 180–200 cm in fully extended position.
7. Height Adjustment Range: Approximately 55–85 cm.
8. Trendelenburg Angle: Minimum 12°–15°.
9. Backrest Angle: Adjustable from 0° up to at least 75° or more.
10. Leg Rest: Independently adjustable or synchronized movement with backrest.
Control System
1. Handheld or integrated control panel for patient use.
2. Separate lockable nurse control panel for clinical staff.
3. Memory function for at least 2–3 preset positions.
4. One-touch CPR emergency function returning the chair to a fully flat position.
Safety Features
1. Anti-pinch protection system for all moving parts.
2. Emergency stop button.
3. Lockable medical-grade castor wheels.
4. Stable anti-tip base design during movement and adjustment.
5. Compliance with IEC medical electrical safety standards.
Comfort Features
1. Adjustable ergonomic headrest.
2. Removable or swing-away armrests.
3. High-density medical-grade cushioning for pressure relief.
4. Anti-decubitus, breathable, waterproof upholstery material.
Electrical Requirements
1. Power supply: 220–240V, 50/60 Hz.
2. Low-noise electric motor system.
3. Manual emergency override system for power failure situations.
Accessories
1. Adjustable IV pole (infusion stand).
2. Side rails for accessory mounting.
3. Optional paper roll holder.
4. Optional storage basket.
Materials & Finish
1. Frame: Powder-coated steel or high-grade aluminum alloy.
2. Upholstery: Medical-grade PU leather or equivalent material.
3. All components shall be corrosion-resistant and compatible with hospital disinfectants.
Certifications & Standards
1. CE Mark (EU MDR 2017/745).
2. ISO 13485 – Medical Device Quality Management System.
3. IEC 60601-1 – Medical electrical safety standard.
4. ISO 10993 – Biocompatibility of patient-contact materials.
5. ISO 14971 – Risk management system (if applicable).
Documentation Requirements
The supplier shall provide the following documents:
1. Original product catalogue with full technical specifications.
2. Declaration of conformity to applicable medical device regulations.
3. CE Certificate or equivalent international approval.
4. ISO 13485 Quality Management System Certificate.
5. IEC 60601-1 compliance certificate.
6. ISO 10993 biocompatibility certificate for patient-contact materials.
7. Certificate of Origin.
8. Factory Acceptance Test (FAT) report.
9. Installation, operation, and maintenance manuals (English, printed and digital).
10. Warranty certificate stating duration and coverage.
11. Preventive maintenance schedule and service plan.
12. Recommended spare parts list (if applicable).
13. Training program for clinical and technical staff.
Spare Parts Availability (Provision of Spare Parts)
1. The supplier shall guara+I173:I178ntee the availability of original spare parts for a minimum period of 5–10 years after installation.
2. All spare parts shall be genuine, manufacturer-approved, and compatible with the supplied equipment.
3. The supplier shall maintain adequate stock of critical spare parts to ensure minimum downtime of the equipment.
4. Delivery time for essential spare parts shall not exceed 7–14 working days from the date of order (or as agreed in the contract).
5. The supplier shall provide a recommended list of critical spare parts for preventive maintenance and emergency repairs.
6. The supplier shall ensure availability of technical support for installation and replacement of spare parts either locally or through authorized service agents
Warranty
1. Minimum2 years warranty against manufacturing defects.
2. Warranty shall include technical support and maintenance services.