## Specs - 6-2026
Technical Datasheet - RFQ - 06-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | Infant Incubator – Technical Specification
1. Functional Description
The infant incubator shall provide a closed, controlled environment for neonatal care by maintaining stable body temperature. Heat is delivered through circulating warm air, ensuring safe and consistent thermal support for the infant.
2. Operational Requirements
Basic infant incubator suitable for general neonatal care
Temperature control in:
Servo (patient) mode
Manual (air) mode
Constructed from transparent, medical-grade materials for visibility and hygiene
Design optimized for ease of use and maintenance
3. System Configuration
Complete incubator unit supplied with essential accessories required for operation
4. Technical Specifications
4.1 Control System
Electronic control system (microprocessor preferred)
Simple digital display for temperature monitoring
Basic alarm functions for safety (high/low temperature)
4.2 Bed and Base
Fixed height or manually adjustable base
Manual bed tilt up to ±8° (head-up and head-down)
Simple locking mechanism
4.3 Temperature Control
Servo (patient) mode: 35°C – 37°C
Air (manual) mode: 20°C – 37°C
Accuracy: ±0.3°C to ±0.5°C
Uniform air circulation within the chamber
Air velocity within safe limits
4.4 Humidity Control
Basic or optional humidity provision (non-servo controlled)
Approximate humidity range: 30% – 70%
Manual adjustment acceptable
4.5 Oxygen Control
No integrated oxygen control system required
Provision for connection to an external oxygen source
4.6 Design and Materials
Single-wall or basic double-wall canopy
Minimum four (4) access hand ports
Additional ports for tubing and monitoring leads
Medical-grade foam mattress (washable and waterproof)
Provision for IV pole
4.7 Mobility
Four swivel castors
At least two wheels with locking mechanism
4.8 Physical Characteristics
Noise level: ≤ 55 dB
Robust construction suitable for continuous use
4.9 Monitoring Features
Basic temperature display
No requirement for:
Integrated weighing scale
X-ray cassette tray
Advanced data trending or monitoring systems
5. Accessories, Spares, and Consumables
One (1) additional mattress
One (1) additional temperature sensor
Essential accessories required for operation
6. Operating Environment
Designed to operate under local environmental conditions, including:
High ambient temperatures
Humidity variations
Dust exposure
Power supply:
220–240 VAC, 50 Hz
Suitable UPS with maintenance free batteries, voltage regulation and spike protection for minimum 30 min. back-up shall be supplied with the system.
Standard power cable and plug
7. Standards and Safety Requirements
Compliance with at least one of the following:
ISO 13485
CE certification
Compliance with:
IEC 60601-2-19 (Particular requirements for infant incubators)
8. User Training
Supplier shall provide basic training covering:
Operation of the device
Routine care and cleaning
Basic safety procedures
9. Warranty
Minimum one (2) years warranty covering manufacturing defects and functionality
10. Maintenance During Warranty
Maintenance provided upon request during the warranty period
Includes corrective (breakdown) maintenance
11. Documentation
User (operating) manual in English (Arabic optional)
Basic service or maintenance instructions
List of essential spare parts (optional)
2 | Phototherapy Machine – Technical Specification
1. Functional Description
The phototherapy unit shall provide effective light therapy for the treatment of neonatal jaundice by delivering light in the appropriate wavelength range to reduce serum bilirubin levels safely.
2. Operational Requirements
Suitable for use in neonatal units and general wards
Easy to operate and maintain
Compatible with infant incubators and open care systems
Designed for continuous operation
3. System Configuration
Complete phototherapy unit with light source, stand, and essential accessories
4. Technical Specifications
4.1 Light Source
LED-based phototherapy unit (preferred for low power consumption and long life)
Wavelength range: 430 – 490 nm (effective blue light spectrum)
Irradiance level:
≥ 20 µW/cm²/nm at mattress level (minimum acceptable)
Uniform light distribution over treatment area
4.2 Control System
Simple control panel with ON/OFF switch
Basic timer function (optional)
Indicator for power status
4.3 Height Adjustment
Manually adjustable height
Adjustable angle for optimal positioning over infant
4.4 Design and Structure
Lightweight and compact design
Mounted on mobile stand or trolley
Durable and easy-to-clean surface
Overhead type (recommended for cost efficiency)
4.5 Mobility
Four castors for easy movement
At least two lockable wheels
4.6 Safety Features
Low heat emission (LED technology)
Overheating protection (basic)
Eye protection for infant (eye shields included)
Electrical safety compliant design
4.7 Noise Level
Silent operation (no fan or minimal noise)
5. Accessories
Eye protectors (reusable or disposable)
One spare set of LED module
Basic user accessories
6. Operating Environment
Suitable for use in:
High ambient temperatures
Humid and dusty environments
Power supply:
220–240 VAC, 50 Hz
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Compliance with:
IEC 60601-1 (general medical electrical safety)
IEC 60601-2-50 or relevant phototherapy standards
8. User Training
Basic training on:
Operation
Positioning
Safety precautions
9. Warranty
Minimum 2-years warranty
10. Maintenance During Warranty
On-demand maintenance during warranty period
11. Documentation
User manual (English, Arabic)
Basic maintenance instructions
3 | Fluorescence Immunoassay Analyzer – Technical Specification
1. Functional Description
The Fluorescence Immunoassay Analyzer shall be a compact, bench-top device used for the quantitative or qualitative detection of analytes in human samples (e.g., blood, serum, plasma, or urine) using fluorescence-based immunoassay technology.
2. Operational Requirements
Suitable for small laboratories, clinics, and point-of-care settings
Easy to operate with minimal training
Fast turnaround time for test results
Capable of continuous or batch sample testing
3. System Configuration
Complete analyzer unit
Compatible test kits (assay-specific)
Basic accessories required for operation
4. Technical Specifications
4.1 Measurement Principle
Fluorescence immunoassay (FIA) technology
Capable of detecting fluorescence intensity and converting it into analyte concentration
4.2 Test Capability
Supports common tests such as:
Infectious diseases (e.g., malaria, dengue)
Cardiac markers (e.g., Troponin, CK-MB)
Hormones (e.g., TSH, HCG)
Open or semi-closed system preferred (compatible with multiple kits if possible)
4.3 Throughput
Single test or small batch processing
Test time: approximately 5–15 minutes per sample
4.4 Sample Types
Whole blood, serum, plasma, or urine
4.5 Detection System
Built-in fluorescence detection module
Automatic calibration (or calibration via test kit)
4.6 Control System
Microprocessor-controlled
Simple user interface with keypad or touchscreen
Basic display for results
4.7 Data Management
Internal memory for storing results
Basic result review and recall
USB port or similar for data export (optional)
4.8 Calibration and Quality Control
Calibration via test kit or internal standard
Basic quality control functionality
4.9 Design and Construction
Compact, lightweight bench-top design
Durable outer casing
Easy to clean and maintain
4.10 Power Requirements
220–240 VAC, 50 Hz
Low power consumption
5. Accessories, Spares, and Consumables
Starter kit of test cartridges/reagents
Calibration materials (if required)
Basic consumables (pipette tips, sample holders if applicable)
Attach the cost per test in details (Euro) Mandatory
Attach the cost of each kits and reagent and consumbles (Euro) Mandatory
Attach the cost of all spare parts (Euro) Mandatory
6. Operating Environment
Suitable for:
High temperature environments
Humidity variations
Dust exposure
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Electrical safety compliance:
IEC 61010-1 (laboratory equipment safety)
8. User Training
Basic training covering:
Operation
Test procedures
Basic troubleshooting
9. Warranty
Minimum 2-years warranty
10. Maintenance During Warranty
On-demand maintenance during warranty period
11. Documentation
User manual (English, Arabic optional)
Basic service/maintenance instructions
4 | Oxygen Concentrator (5 LPM) – Technical Specification
1. Functional Description
The oxygen concentrator shall provide a continuous supply of medical-grade oxygen by extracting oxygen from ambient air using pressure swing adsorption (PSA) technology. The unit shall be suitable for use in hospitals, clinics, and home care settings.
2. Operational Requirements
Continuous oxygen supply up to 5 liters per minute (LPM)
Simple operation with minimal user training
Reliable performance under varying environmental conditions
Low maintenance requirements
3. System Configuration
Complete oxygen concentrator unit
Integrated compressor, sieve beds, and control system
Supplied with essential accessories
4. Technical Specifications
4.1 Oxygen Output
Flow rate: 0.5 – 5 LPM adjustable
Oxygen concentration:
93% at 2–5 LPM ±3%
Continuous operation for 10-12 hours .
4.2 Technology
Pressure Swing Adsorption (PSA) technology
Oil-free compressor
4.3 Control System
Simple control panel with:
Flow control knob
Power ON/OFF switch
Basic indicators:
Power status
Oxygen purity (basic indicator preferred)
4.4 Safety Features
High/low pressure alarm
Low oxygen concentration alarm (recommended but optional in economic version)
Power failure alarm
Overheating protection
4.5 Noise Level
≤ 55 dB
4.6 Design and Construction
Compact and portable design
Durable casing suitable for continuous use
Integrated handle for movement
4.7 Mobility
Two or four castor wheels for easy movement
4.8 Filters
Dust air filter (washable and reusable)
4.9 Power Requirements
220–240 VAC, 50 Hz
Low power consumption
Power cable with standard plug
5. Accessories, Spares, and Consumables
Nasal cannula (adult/pediatric)
Humidifier bottle
Intake air filter for compressor qty 12pcs.
Spare air filter (at least one)
User accessories for basic operation
6. Operating Environment
Designed for use in:
High ambient temperatures (up to ~40°C recommended)
Humidity variations
Dusty environments
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Electrical safety:
IEC 60601-1
8. User Training
Basic training covering:
Operation
Cleaning of filters
Basic troubleshooting
9. Warranty
Minimum 2-years warranty covering defects and function
Preventive and corrective maintenance on request
Basic troubleshooting guide provided
11. Documentation
User manual (English, Arabic optional)
Maintenance manual
List of spare parts and accessories
5 | Oxygen Concentrator (10 LPM) – Technical Specification
1. Functional Description
The oxygen concentrator shall provide a continuous supply of medical-grade oxygen by filtering and concentrating oxygen from ambient air using Pressure Swing Adsorption (PSA) technology. It shall be suitable for clinical and home use.
2. Operational Requirements
Continuous oxygen delivery up to 10 liters per minute (LPM)
Designed for reliable operation in low-resource settings
Easy to operate and maintain
Suitable for continuous (16/7) use
3. System Configuration
Complete oxygen concentrator unit
Built-in compressor, filters, and sieve beds
Standard accessories for operation
4. Technical Specifications
4.1 Oxygen Output
Flow rate: 1 – 10 LPM (adjustable)
Oxygen concentration:
93% at 2–5 LPM ±3%
4.2 Technology
Pressure Swing Adsorption (PSA) technology
Dual sieve bed system
4.3 Control System
Simple control panel
Flow meter with manual adjustment knob
Power ON/OFF switch
Basic indicator lights
4.4 Alarms and Safety Features
Power failure alarm
Pressure fault alarm
4.5 Filtration System
Dust filter (washable/reusable)
Intake air filter suitable for dusty environments qty:12pcs.
4.6 Design and Construction
Compact and robust cabinet
Durable outer casing
Designed for easy maintenance
Integrated humidifier bottle holder
4.7 Mobility
Four wheels for easy movement
Lightweight design with handle
4.8 Noise Level
≤ 60 dB
4.9 Operating Conditions
Ambient temperature: up to 40°C or higher
Humidity: suitable for tropical environments
Designed for dusty conditions
4.10 Power Requirements
220–240 VAC, 50 Hz
Power consumption: ≤ 600–800 W
5. Accessories, Spares, and Consumables
Humidifier bottle
Nasal cannula and oxygen tubing
Spare air filters
User-replaceable components where possible
6. Operating Environment
Suitable for use in:
High temperature areas
High humidity
Dusty conditions (e.g., Sudan and similar environments)
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Compliance with:
IEC 60601-1 (medical electrical safety)
8. User Training
Basic training covering:
Cleaning of filters
Operation
Flow adjustment
Maintenance and troubleshooting
9. Warranty
Minimum 2-years warranty covering defects and function
Preventive and corrective maintenance on request
Basic troubleshooting guide provided
11. Documentation
User manual (English, Arabic optional)
Maintenance manual
List of spare parts and accessories
6 | Hemoglobin Meter – Technical Specification
1. Functional Description
The hemoglobin meter shall be a portable device used for the quantitative measurement of hemoglobin (Hb) concentration in human blood, supporting rapid screening and diagnosis of anemia.
2. Operational Requirements
Suitable for point-of-care use (clinics, wards, outreach settings)
Easy to operate with minimal training
Rapid test results
Battery-operated for use in low-resource settings
3. System Configuration
Portable hemoglobin meter
Compatible test strips or microcuvettes
Basic accessories required for operation
4. Technical Specifications
4.1 Measurement Principle
Photometric or reflectance method for hemoglobin determination
4.2 Measurement Range
Approximate range: 0 – 25 g/dL
4.3 Accuracy
±1.0 g/dL or better
4.4 Sample Type and Volume
Capillary or venous whole blood
Sample volume: ≤ 20 µL
4.5 Test Time
Results available within 10–60 seconds
4.6 Display and Control
Digital display (LCD or similar)
Simple operation with minimal buttons
4.7 Data Management
Internal memory for storing results (basic capacity)
USB or data transfer capability
4.8 Power Supply
Rechargeable battery or replaceable batteries
Optional AC adapter (220–240 VAC, 50 Hz)
4.9 Design and Construction
Lightweight and portable
Durable casing suitable for field use
Easy to clean
5. Accessories, Spares, and Consumables
Starter supply of test strips or microcuvettes
Lancets and capillary collection accessories (if applicable)
Carrying case
6. Operating Environment
Suitable for:
High ambient temperatures
Humidity variations
Dust exposure
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Electrical safety compliance:
IEC 61010-1
8. User Training
Basic training on
8. User Training
Basic training covering:
Sample collection
Operation and result interpretation
Cleaning and maintenance
9. Warranty
Minimum 2-years warranty covering device function and defects
10. Maintenance During Warranty
On-demand maintenance service
Basic troubleshooting guide provided
11. Documentation
User manual (English, Arabic optional)
Basic maintenance and troubleshooting instructions
List of consumables and accessories
7 | Infant Warmer (Radiant Warmer) – Technical Specification
1. Functional Description
The infant warmer shall provide open, radiant heat to maintain a stable body temperature for neonates during resuscitation, postnatal care, and neonatal procedures.
2. Operational Requirements
Open care system radiant warmer
Suitable for delivery rooms, neonatal units, and emergency care
Rapid heat-up capability
Simple and reliable operation with minimal training
3. System Configuration
Complete radiant infant warmer unit
Mounted on mobile stand with mattress platform
Supplied with essential accessories
4. Technical Specifications
4.1 Heating System
Infrared radiant heating system
Quartz or ceramic heating element (cost-effective option)
Manual or basic servo temperature control
Warm-up time: fast stabilization (≤ 5 minutes recommended)
4.2 Temperature Control
Skin (servo) mode: 34°C – 37°C
Manual mode available
Accuracy: ±0.5°C
Over-temperature protection included
4.3 Control System
Microprocessor-based or basic electronic control
Digital display for temperature
Simple control panel with key functions only
Basic alarm system (high temperature alarm mandatory)
4.4 Mattress and Bed
Flat or slightly contoured infant mattress
Removable and washable cover
Optional tilting bed (manual adjustment preferred)
4.5 Structure and Design
Open-access design for easy clinical procedures
Mounted on sturdy mobile stand
Durable metal frame with anti-rust coating
4.6 Safety Features
Overheat protection system
Heater failure alarm
Skin probe failure alarm (if servo mode included)
Electrical safety protection
4.7 Mobility
Four castors for easy movement
At least two lockable wheels
4.8 Power Requirements
220–240 VAC, 50 Hz
Low power consumption
5. Accessories
Skin temperature probe (if servo mode included)
One spare heater element (recommended)
Basic infant mattress
Power cable and standard accessories
6. Operating Environment
Suitable for tropical and subtropical conditions:
High ambient temperature
High humidity
Dusty environments
7. Standards and Safety Requirements
Compliance with at least one:
CE certification
ISO 13485
Electrical safety compliance:
IEC 60601-1
IEC 60601-2-21 (Infant radiant warmers)
8. User Training
Basic training covering:
Operation
Temperature control
Safety procedures
9. Warranty
Minimum 2-years warranty
10. Maintenance During Warranty
On-demand maintenance service
Basic preventive maintenance support
11. Documentation
User manual (English, Arabic optional)
Basic service and maintenance instructions
8 | 1.Mobile X-Ray Machine-Technical Specifications
A microprocessor controlled mobile x-ray imaging unit is to include but not limited to the following:
Microprocessor controlled x-ray generator, > 7.5 KW
It should have a digital display of mAs and kV and an electronic timer.
KV range:40kV to 110 kV
mA range: 300 mA or more. Please specify mA and seconds separately and not mAs alone.
Shortest exposure time: I ms.
X-Ray Tube: .
Output should match the output of the generator,
Must have a rotating anode with at least: 2500 rpm and focal spot size should be less than Imm.
Rotating anode tube system, with focal spots , 0.6/1.2 mm of large heat storage capacity, > 90,000 HU
Collimator- Manually adjustable multileaf collimator, rotatable ±90°
The exposure release switch should be detachable with a cord of at least 5 meters
Remote control operating distance > 10 metres
Remote control operating radius- 180 deg
Lightweight manual driven unit, with braking system.
Small source image distance, please specify.
Direct 220-240v/50Hz single phase power line connection with built in line voltage compensation.
Grid( Ratio 6:1) of following sizes should be provided (1) each 12”x15” & 10”x12”
2.User Training:
2.1The Supplier shall conduct user training for this equipment to enable operators to use the equipment properly. The training shall include the use of all operational functions of the equipment, as well as routine checks and maintenance expected by users.
3.Warranty
3.1Comprehensive warranty for 2 years from acceptance.
4.Maintenance Service During Warranty Period
4.1During warranty period supplier must ensure preventive maintenance & corrective/breakdown maintenance whenever required.
5.Installation and Commissioning
5.1The equipment to be installed by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
6.Documentation
6.1User (Operating) manual in English.
6.2Service (Technical / Maintenance) manual in English.
6.3List of important spare parts and accessories with their part numbers and costing.
9 | C-Arm -Technical Specifications
1.Description of Function
1.1This equipment is used in orthopedic fractures for imaging of bone pathology or fractures on display monitor during operation/ reduction of fractures.
2.Operational Requirements For continuous fluoroscopy, image storage and retrieval the system offered shall be a general fluoroscopy/radiology system
3.System Configuration
3.1X-ray C-Arm Mobile with complete accessories
4.Technical Specifications
4.1X-Ray Generator
X-Ray Generator Microprocessor based, high frequency inverter generator
Generator Output: 3.5kW or more
Fluoroscopic/ Radiographic KV range Lower limit shall not exceed 40 KV Higher limit shall not be less than 110 KV
Fluoroscopic mA range Lower limit shall be ~0.1 mA Upper limit shall be ~9 mA
4.2X-Ray Tube
Rotating anode type
Small focal spot, shall not be more than 0.3 mm
Large focal spot, shall not be more than 0.6 mm
Nominal voltage: 110 KV
Anode heat storage capacity not less than 70 KHU
Inherent filtration should be at least 0.5 mm Al eq
4.3Collimator
Operator controlled automatic collimation
4.4 C-Arm · Focus - I.I. Distance shall be at least 100 cm
· Depth shall be ~ 60cm
· Horizontal travel at least 200mm
· Vertical travel at least 430 mm
· Orbital movement shall be ~120°
· Swivel range shall be ~12°
· Rotation about horizontal axis shall be more than +/-180°
4.5Image Intensifier:
· Triple Mode 9” with direct coupling with camera
· Shall be at least 52 lp/ cm
· Noise reduction, scattered light trap for high contrast dynamics
· CCD camera technology with ABC and AGC control
4.6TV Monitor
· One unit LCD monitor side by side for live and reference image
· Shall be at least 32” with automatic brightness control
4.7Image rotation
· Shall be at least 625 scanning lines at 50 Hz
· Trolley for one display screens and with the alphanumeric keyboard included
· High resolution
4.8 Imaging Modes
· Fluoroscopy mode shall have the following facilities:
· Continuous fluoroscopy with last image hold
· Last image holds with at least two frames image memory
· Continuous fluoroscopy with image acquisition rate: about 20 frame/second.
· Hard disk with image storage capacity of at least 30000 images
· RAM Memory of 256 images
· Mosaic display of 16 images · Zoom (x 2)
· Come with CD/DVD/RW drive
5.System Configuration Accessories, spares and consumables
5.1 Sterilizable textile cover and clips, for the X-ray tube and the Cassette holder for 24 x 30 cm
5.2 All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer.
6. Operating Environment
6.1The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature, Humidity, etc. for Sudan.
6.2 Should work on 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs and sockets.
Standards and Safety Requirements
The unit offered shall be certified to meeting the relevant requirements of TUV, CE mark (MDD), FDA and/ or any equivalent quality and safety standards.
8. User Training
8.1On site operational training till the familiarity of the system and satisfaction of end user shall be provided.
9.Warranty
9.1 Comprehensive warranty for 2 years.
10 Maintenance Service During Warranty Period
10.1Preventive & Corrective Maintenance
11. Installation, Inspections and Commissioning
11.1Supplier must accomplish proper installation and commissioning of the equipment on site
12.Documentation
12.1 User (Operating) manual in English
12.2Service (Technical / Maintenance) manual in English
10 | ANESTHESIA MACHINE – TECHNICAL SPECIFICATION (ECONOMIC VERSION)
1. General Requirements
1. Anesthesia machine suitable for safe delivery of medical gases and volatile anesthetic agents.
2. Suitable for adult and pediatric patients; neonatal use preferred.
3. Basic electronic control system with digital display.
4. Mounted on trolley with four swivel castors (minimum two lockable).
5. Compact, durable, and easy-to-clean design suitable for operating room use.
6. Integrated breathing system and ventilator.
7. Compatible with standard patient monitoring systems.
2. Gas Supply & Flow Control
Gas inputs: O₂, N₂O
Cylinder yokes for O₂ and N₂O with PIN-index safety system
Mechanical or electronic flowmeters
Oxygen flow range: 0–10 L/min
N₂O flow range: 0–10 L/min
Oxygen flush valve included
Hypoxic guard system to prevent delivery of low oxygen mixtures
Backup battery support: minimum 30 minutes
3. Breathing Circuit & CO₂ Absorption
Reusable breathing circuits for adult and pediatric patients
Pediatric circuit included
CO₂ absorber canister with soda lime capacity ≥1.5 kg
Adjustable Pressure Limiting (APL) valve
Manual and automatic ventilation switching
Components detachable for cleaning and sterilization
4. Vaporizers
Minimum two (2) agent-specific vaporizers included:
Isoflurane
Sevoflurane
Temperature and flow compensated
Interlock safety system to prevent simultaneous
Easy and secure filling system with leak prevention
Maintenance-free or low-maintenance design preferred
Mounted securely on the anesthesia machine with locking mechanism
5. Ventilator
Electrically powered ventilator with battery backup
Ventilation modes:
Manual
Volume Control Ventilation (VCV)
Pressure Control Ventilation (PCV)
Spontaneous mode
Tidal volume range:
20–1500 mL
Respiratory rate:
4–60 bpm
PEEP:
0–20 cmH₂O
Pressure monitoring included
6. Monitoring Parameters
Display of:
Airway pressure
Tidal volume
Respiratory rate
Minute ventilation
FiO₂
Optional:
EtCO₂ monitoring
Agent monitoring
7. Alarm & Safety System
Audio-visual alarms for:
Low oxygen pressure
Power failure
High airway pressure
Disconnection alarm
Oxygen failure protection device
Emergency oxygen bypass
8. Display & User Interface
Basic color display or LCD screen
Display of ventilation parameters and alarms
Simple keypad or control knobs for adjustment
9. Accessories & Consumables
Adult breathing circuit
Pediatric breathing circuit
Breathing bag
Gas hoses
Oxygen sensor
One spare set of filters and seals
CO₂ absorbent starter pack
10. Operating & Storage Conditions
Suitable for Sudan climate conditions:
Operating temperature:
20°C – 40°C
Relative humidity:
20% – 80%
Storage temperature:
5°C – 45°C
Power supply:
220–240 VAC, 50 Hz
11. Standards & Safety Compliance
The system shall comply with:
ISO 13485 or CE certification or FDA Or equivelent
IEC 60601-1
IEC 60601-2-13
12. Training
Supplier shall provide basic training covering:
System operation
Safety procedures
Maintenance and Troublshooting .
Cleaning and routine maintenance
13. Warranty & Maintenance
Minimum 2-years warranty
Preventive and corrective maintenance during warranty period
Availability of spare parts for minimum 10 years .
14. Documentation
Supplier shall provide:
Service manual (English and Arabic)+A759:D761
Service manual (English)
Spare parts list (Euro)
Calibration and inspection certificates
11 | ICU VENTILATOR – TECHNICAL SPECIFICATION
(Adult and Pediatric Use)
1. General Requirements
1. ICU ventilator suitable for invasive and non-invasive ventilation in intensive care units.
2. Suitable for adult and pediatric patients.
3. Compact, microprocessor-controlled ventilator with user-friendly operation.
4. Designed for continuous operation in hospital environments.
5. Mounted on trolley with swivel castors (minimum two lockable).
6. Integrated turbine and central gas supply compatible.
2. Ventilation Modes
The ventilator shall provide at minimum the following modes:
Volume Control Ventilation (VCV)
Pressure Control Ventilation (PCV)
SIMV (Volume and/or Pressure)
CPAP
Pressure Support Ventilation (PSV)
Spontaneous Breathing Mode
Non-Invasive Ventilation (NIV)
3. Ventilation Parameters
Tidal volume range:
50–2000 mL
Respiratory rate:
1–80 breaths/min
PEEP:
0–30 cmH₂O
Inspiratory pressure:
5–80 cmH₂O
FiO₂ range:
21%–100%
I:E ratio adjustable
Trigger sensitivity adjustable
4. Monitoring Parameters
Real-time display of:
Tidal volume
Respiratory rate
Minute ventilation
Peak airway pressure
Mean airway pressure
PEEP
FiO₂
Flow and pressure waveforms
Optional:
EtCO₂ monitoring
Lung compliance and resistance
5. Alarm & Safety System
Audio-visual alarms for:
High/low airway pressure
Apnea
Power failure
Gas supply failure
Low tidal volume
High respiratory rate
Disconnection alarm
Additional safety features:
Oxygen failure protection
Internal self-test function
Emergency ventilation mode
6. Display & User Interface
Color LCD or TFT display
Minimum screen size: 8 inches
Real-time waveform display
Simple menu navigation and parameter adjustment
Display of alarm messages and system status
7. Gas Supply
Compatible with:
Central oxygen and Air supply
Built-in turbine system for operation without compressed air source
8. Power Supply & Battery
Power supply:
220–240 VAC, 50 Hz
Internal rechargeable battery backup:
Minimum 60 minutes operation
Battery charging indicator included
9. Accessories & Consumables
The ventilator shall be supplied with:
Adult breathing circuit
Pediatric breathing circuit
Humidifier chamber (basic)
Reusable flow sensor
Test lung
Oxygen hose and air hose
Power cable
Basic filters and consumables
10. Operating & Storage Conditions
Suitable for Sudan climate conditions:
Operating temperature:
10°C – 40°C
Relative humidity:
20% – 85%
Storage temperature:
5°C – 45°C
11. Standards & Safety Compliance
The ventilator shall comply with:
ISO 13485 or CE certification
IEC 60601-1
IEC 60601-1-2
ISO 80601-2-12 (Critical care ventilators)
12. Training
Supplier shall provide user training covering:
Device operation
Ventilation modes
Alarm management
Maintenance and Troublshooting
Cleaning and basic maintenance
13. Warranty & Maintenance
Minimum 2-years warranty
Preventive and corrective maintenance during warranty period
Spare parts availability for minimum 10 years
14. Documentation
Supplier shall provide:
User manual (English)
Service manual (English)
Spare parts list
Calibration and inspection certificates
12 | WHEELCHAIR – TECHNICAL SPECIFICATION
1. General Requirements
1. Manual wheelchair suitable for indoor and outdoor patient transportation.
2. Designed for adult patient use in hospitals, clinics, and healthcare facilities.
3. Lightweight, durable, and easy-to-clean construction.
4. Foldable design for easy storage and transport.
5. Comfortable seating with safe and stable operation.
2. Frame & Structure
Strong steel or aluminum frame
Foldable frame design
Anti-rust coated finish
Reinforced cross-brace structure for stability
Maximum patient weight capacity:
Minimum 120 kg
3. Seat & Backrest
Seat width:
Approximately 45–50 cm
Padded seat and backrest
Waterproof and washable upholstery material
Comfortable armrests fixed or detachable
4. Wheels & Mobility
Rear wheels:
Large solid or pneumatic tires
Approximately 24-inch diameter
Front caster wheels:
Approximately 6–8 inches
Smooth rolling performance
Hand rims for self-propulsion
5. Footrests & Armrests
Swing-away or detachable footrests
Non-slip footplates
Fixed or detachable armrests
Comfortable support for patient transfer
6. Braking System
Manual parking brakes on rear wheels
Easy-to-operate locking mechanism
Stable braking during patient transfer
7. Safety Features
Anti-tip stability design preferred
Rounded edges for patient safety
Stable frame construction
Non-slip hand grips
8. Physical Characteristics
Lightweight and easy to maneuver
Foldable for transportation and storage
Easy cleaning and maintenance
9. Accessories
The wheelchair shall include:
Footrests
Armrests
Seat cushion (basic)
User manual
Optional:
IV pole holder
Safety belt
Anti-tip wheels
10. Operating Conditions
Suitable for hospital and general healthcare environments, including:
High temperature conditions
Humid environments
Frequent daily use
11. Standards & Safety Compliance
The wheelchair shall comply with:
ISO 13485 or CE certification preferred
ISO 7176 (Wheelchair standards) preferred
12. Warranty
Minimum 2-years warranty against manufacturing defects
13. Documentation
Supplier shall provide:
User manual (English)
Basic maintenance instructions
Spare parts list (optional)
13 | Dental Extracting Set – Adult- TECHNICAL SPECIFICATION
1. General Description
Supply of an Adult Dental Extracting Set suitable for use in dental clinics, hospitals, and oral surgery departments. The set shall be manufactured from high-quality medical-grade stainless steel, reusable, autoclavable, corrosion-resistant, and supplied with a comprehensive warranty of not less than two (2) years against manufacturing defects and rust.
2. Technical Specifications
Product Name Dental Extracting Set – Adult
Intended Use Adult tooth extraction procedures
Number of Pieces Minimum 10–12 pieces
Material High-quality medical-grade stainless steel
Stainless Steel Grade AISI 420 or equivalent
Corrosion Resistance Rust-resistant and corrosion-resistant
Sterilization Fully autoclavable up to 134–135°C
Surface Finish Highly polished satin or mirror finish
Handle Design Ergonomic, non-slip handle
Beak Design Anatomically contoured and serrated beaks
Reusability Fully reusable
Storage Case Stainless steel or aluminum sterilization case included
Quality Standards CE Marked and/or ISO 13485 and/or FDA compliant
3. Minimum Set Contents
The set shall include, at minimum, the following extraction forceps or equivalent:
1. Upper Anterior Forceps No.1
2. Upper Premolar Forceps No.7
3. Lower Premolar Forceps No.13
4. Upper Molar Right Forceps No.17
5. Upper Molar Left Forceps No.18
6. Upper Root Forceps No.51
7. Upper Third Molar Forceps No.67
8. Lower Molar Forceps No.73
9. Lower Root Forceps No.74
10. Lower Molar Cowhorn Forceps No.86
Equivalent international numbering systems may be accepted.
4. Quality and Compliance Requirements
The offered product shall comply with the following requirements:
Manufactured in accordance with ISO 13485 and/or ISO 9001 standards.
CE certified and/or FDA compliant.
Resistant to corrosion, staining, and repeated sterilization cycles.
Smooth joint movement without excessive looseness.
Easy to clean and disinfect.
Suitable for repeated clinical use.
Laser marking of instrument identification is preferred.
Ultrasonic cleaned and passivated instruments are preferred.
5. Warranty and After-Sales Service
The supplier shall provide:
Minimum two (2) years warranty against:
Manufacturing defects
Joint failure
Corrosion and rust under normal use
Replacement of defective instruments during the warranty period.
Availability of spare or replacement instruments when required.
6. Delivery Requirements
Delivery period shall not exceed 30–45 days from the date of purchase order.
All items must be supplied in original manufacturer packaging.
The purchaser reserves the right to reject any item not complying with the required specifications.
7. Commercial and Competitive Requirements
To ensure competitive pricing in public tendering:
Bidders shall submit the following documents with their offers:
Product catalog and technical datasheet
Certificate of origin
CE certificate and/or ISO certification
Warranty statement
Manufacturer or authorized distributor authorization
14 | Dental Extracting Set – Pediatric-TECHNICAL SPECIFICATIONS
1. General Description
Supply of a Pediatric Dental Extracting Set suitable for use in dental clinics, hospitals, and oral surgery departments for children. The set shall be manufactured from high-quality medical-grade stainless steel, reusable, autoclavable, corrosion-resistant, and supplied with a comprehensive warranty of not less than two (2) years against manufacturing defects and rust.
2. Technical Specifications
Product Name :Dental Extracting Set – Pediatric
Intended Use :Pediatric tooth extraction procedures
Number of Pieces: Minimum 6–10 pieces
Material: High-quality medical-grade stainless steel
Stainless Steel Grade: AISI 420 or equivalent
Corrosion Resistance : Rust-resistant and corrosion-resistant
Sterilization: Fully autoclavable up to 134–135°C
Surface Finish :Highly polished satin or mirror finish
Handle Design: Ergonomic, lightweight, non-slip handle
Beak Design Pediatric anatomical design with fine serrated beaks
Reusability: Fully reusable
Storage Case :Stainless steel or aluminum sterilization case included
Quality Standards:CE Marked and/or ISO 13485 and/or FDA compliant
Product Condition:Brand new, unused, non-refurbished
3. Minimum Set Contents
The set shall include, at minimum, the following pediatric extraction forceps or equivalent:
1. Pediatric Upper Anterior Forceps
2. Pediatric Upper Molar Forceps
3. Pediatric Upper Root Forceps
4. Pediatric Lower Anterior Forceps
5. Pediatric Lower Molar Forceps
6. Pediatric Lower Root Forceps
Additional pediatric extraction forceps may be included depending on manufacturer configuration.
Equivalent international numbering systems may be accepted.
4. Quality and Compliance Requirements
The offered product shall comply with the following requirements:
Manufactured in accordance with ISO 13485 and/or ISO 9001 standards.
CE certified and/or FDA compliant.
Resistant to corrosion, staining, and repeated sterilization cycles.
Smooth joint movement without excessive looseness.
Fine pediatric design suitable for primary teeth extraction.
Easy to clean and disinfect.
Suitable for repeated clinical use.
Laser marking of instrument identification is preferred.
Ultrasonic cleaned and passivated instruments are preferred.
5. Warranty and After-Sales Service
The supplier shall provide:
Minimum two (2) years warranty against:
Manufacturing defects
Joint failure
Corrosion and rust under normal use
Replacement of defective instruments during the warranty period.
Availability of spare or replacement instruments when required.
6. Delivery Requirements
6.1. Required Submission Documents
Bidders shall submit the following documents with their offers:
Product catalog and technical datasheet
Certificate of origin
CE certificate and/or ISO certification
Warranty statement
Manufacturer or authorized distributor authorization (if applicable