## Sheet1
Technical Datasheet - RFQ - 07-2026
No | TECHNICAL SPECIFICATION | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | UROLOGY ENDOSCOPIC INSTRUMENT SET
1. GENERAL DESCRIPTION
The supplier shall provide a complete set of reusable rigid urology endoscopic surgical instruments intended for diagnostic and therapeutic procedures in the lower and upper urinary tract.
The system shall be fully compatible with standard rigid endoscopy camera systems, light sources, and monitors already available in the hospital (endoscopy tower excluded).
The set shall support the following procedures:
Cystoscopy
Transurethral resection (TURP / TURBT)
Visual internal urethrotomy (VIU)
Ureteroscopy (URS)
Stone management and lithotripsy
Bladder stone crushing (Punch technique)
All instruments shall be of medical grade, autoclavable, reusable, and manufactured from high-quality stainless steel suitable for repeated sterilization.
2. CYSTOSCOPY INSTRUMENT SET
The set shall include:
Rigid cystoscope sheaths in various sizes (approx. 18Fr–26Fr)
Obturators for each sheath size
Working elements compatible with monopolar and bipolar systems
Telescope holders/adapters compatible with standard rigid endoscopes
Irrigation stopcocks and connectors
3. TRANSURETHRAL RESECTION (TUR) SET WITH DIATHERMY
The system shall include a complete resectoscopy set suitable for monopolar and/or bipolar electrosurgical procedures.
Components:
Continuous flow resectoscope sheaths (approx. 24Fr–26Fr)
Working element compatible with electrosurgical energy
Cutting loops (standard and prostate type)
Coagulation electrodes (roller type)
Vaporization or button electrodes (optional depending on system)
Obturator
Tissue evacuation system (e.g., evacuator syringe type)
Electrosurgical Accessories:
Monopolar and bipolar connection cables
Foot switch with cut and coagulation control
Neutral electrode plate (for monopolar systems)
4. VISUAL INTERNAL URETHROTOMY (VIU) SET
The VIU set shall include:
Optical urethrotome system
Cold knife blade assembly for internal urethral incision
Optical sheath compatible with rigid endoscope system
Telescopes with appropriate viewing angles (e.g., 0°–12°)
Irrigation system compatible with continuous flow
5. URETEROSCOPY (URS) SET
The ureteroscopy set shall include instruments for semi-rigid ureteroscopy procedures.
Components:
Semi-rigid ureteroscopes (various sizes approximately 6Fr–9.8Fr)
Ureteral access sheaths
Hydrophilic and stiff guidewires
Ureteral dilators set
Stone retrieval baskets (Nitinol type)
Ureteral grasping and biopsy forceps
Irrigation system compatible with high flow
6. STONE MANAGEMENT INSTRUMENT SET
The set shall include instruments for stone retrieval and fragmentation support:
Stone grasping forceps (straight and angled types)
Stone retrieval forceps
Nitinol stone baskets (tipless and standard types)
Alligator forceps
Stone fragmentation forceps
Accessories for bladder and ureteric stone extraction
7. BLADDER STONE PUNCH SET
Mechanical stone crushing system suitable for bladder stone management.
Components:
Heavy-duty stone crushing forceps (mechanical punch type)
Straight and curved jaw designs
Compatible cystoscopic working elements
Irrigation-compatible system
Obturator and introducer set
Sterilization tray
8. PNEUMATIC LITHOTRIPSY SYSTEM
A complete pneumatic stone fragmentation system for ureteral and bladder stones.
Components:
Pneumatic lithotripsy generator unit
Foot pedal control system
Reusable handpiece (autoclavable)
Multiple probe sizes (approx. 0.8 mm – 1.2 mm)
Connecting air tubing and accessories
Sterile probe covers and maintenance kit
9. GENERAL REQUIREMENTS
All instruments shall be manufactured from high-grade surgical stainless steel (minimum AISI 316L or equivalent)
All reusable components shall be autoclavable at temperatures up to 134°C
Instruments shall be fully compatible with standard rigid endoscopy systems
Design shall ensure ergonomic handling and safe surgical manipulation
All components shall be supplied with appropriate sterilization trays
10. QUALITY AND STANDARDS
Compliance with international medical device standards (e.g., ISO 13485 or equivalent)
CE-marked or FDA-approved where applicable
Manufacturer shall provide warranty and technical support
Spare parts availability for at least 5 years
11. PACKAGING AND DELIVERY
Instruments shall be supplied in modular, labeled stainless steel sterilization trays
Each set shall be complete and ready for clinical use upon integration with existing endoscopy tower
Training and user manuals shall be provided