## Specs -11-2026
Technical Datasheet - RFQ - 11-2026
No | Technical Specification | Supplier offered specification | Compliant (yes/No) | Remarks / Deviation Details | Supporting Document
1 | ANESTHESIA MACHINE – TECHNICAL SPECIFICATION (ECONOMIC VERSION)
1. General Requirements
1. Anesthesia machine suitable for safe delivery of medical gases and volatile anesthetic agents.
2. Suitable for adult and pediatric patients; neonatal use preferred.
3. Basic electronic control system with digital display.
4. Mounted on trolley with four swivel castors (minimum two lockable).
5. Compact, durable, and easy-to-clean design suitable for operating room use.
6. Integrated breathing system and ventilator.
7. Compatible with standard patient monitoring systems.
2. Gas Supply & Flow Control
Gas inputs: O₂, N₂O
Cylinder yokes for O₂ and N₂O with PIN-index safety system
Mechanical or electronic flowmeters
Oxygen flow range: 0–10 L/min
N₂O flow range: 0–10 L/min
Oxygen flush valve included
Hypoxic guard system to prevent delivery of low oxygen mixtures
Backup battery support: minimum 30 minutes
3. Breathing Circuit & CO₂ Absorption
Reusable breathing circuits for adult and pediatric patients
Pediatric circuit included
CO₂ absorber canister with soda lime capacity ≥1.5 kg
Adjustable Pressure Limiting (APL) valve
Manual and automatic ventilation switching
Components detachable for cleaning and sterilization
4. Vaporizers
Minimum two (2) agent-specific vaporizers included:
Isoflurane
Sevoflurane
Temperature and flow compensated
Interlock safety system to prevent simultaneous
Easy and secure filling system with leak prevention
Maintenance-free or low-maintenance design preferred
Mounted securely on the anesthesia machine with locking mechanism
5. Ventilator
Electrically powered ventilator with battery backup
Ventilation modes:
Manual
Volume Control Ventilation (VCV)
Pressure Control Ventilation (PCV)
Spontaneous mode
Tidal volume range:
20–1500 mL
Respiratory rate:
4–60 bpm
PEEP:
0–20 cmH₂O
Pressure monitoring included
6. Monitoring Parameters
Display of:
Airway pressure
Tidal volume
Respiratory rate
Minute ventilation
FiO₂
Optional:
EtCO₂ monitoring
Agent monitoring
7. Alarm & Safety System
Audio-visual alarms for:
Low oxygen pressure
Power failure
High airway pressure
Disconnection alarm
Oxygen failure protection device
Emergency oxygen bypass
8. Display & User Interface
Basic color display or LCD screen
Display of ventilation parameters and alarms
Simple keypad or control knobs for adjustment
9. Accessories & Consumables
Adult breathing circuit
Pediatric breathing circuit
Breathing bag
Gas hoses
Oxygen sensor
One spare set of filters and seals
CO₂ absorbent starter pack
10. Operating & Storage Conditions
Suitable for Sudan climate conditions:
Operating temperature:
20°C – 40°C
Relative humidity:
20% – 80%
Storage temperature:
5°C – 45°C
Power supply:
220–240 VAC, 50 Hz
11. Standards & Safety Compliance
The system shall comply with:
ISO 13485 or CE certification or FDA Or equivelent
IEC 60601-1
IEC 60601-2-13
12. Training
Supplier shall provide basic training covering:
System operation
Safety procedures
Maintenance and Troublshooting .
Cleaning and routine maintenance
13. Warranty & Maintenance
Minimum 2-years warranty
Preventive and corrective maintenance during warranty period
Availability of spare parts for minimum 10 years .
14. Documentation
Supplier shall provide:
Service manual (English and Arabic)+A759:D761
Service manual (English)
Spare parts list (Euro)
Calibration and inspection certificates
2 | ICU VENTILATOR – TECHNICAL SPECIFICATION
(Adult and Pediatric Use)
1. General Requirements
1. ICU ventilator suitable for invasive and non-invasive ventilation in intensive care units.
2. Suitable for adult and pediatric patients.
3. Compact, microprocessor-controlled ventilator with user-friendly operation.
4. Designed for continuous operation in hospital environments.
5. Mounted on trolley with swivel castors (minimum two lockable).
6. Integrated turbine and central gas supply compatible.
2. Ventilation Modes
The ventilator shall provide at minimum the following modes:
Volume Control Ventilation (VCV)
Pressure Control Ventilation (PCV)
SIMV (Volume and/or Pressure)
CPAP
Pressure Support Ventilation (PSV)
Spontaneous Breathing Mode
Non-Invasive Ventilation (NIV)
3. Ventilation Parameters
Tidal volume range:
50–2000 mL
Respiratory rate:
1–80 breaths/min
PEEP:
0–30 cmH₂O
Inspiratory pressure:
5–80 cmH₂O
FiO₂ range:
21%–100%
I:E ratio adjustable
Trigger sensitivity adjustable
4. Monitoring Parameters
Real-time display of:
Tidal volume
Respiratory rate
Minute ventilation
Peak airway pressure
Mean airway pressure
PEEP
FiO₂
Flow and pressure waveforms
Optional:
EtCO₂ monitoring
Lung compliance and resistance
5. Alarm & Safety System
Audio-visual alarms for:
High/low airway pressure
Apnea
Power failure
Gas supply failure
Low tidal volume
High respiratory rate
Disconnection alarm
Additional safety features:
Oxygen failure protection
Internal self-test function
Emergency ventilation mode
6. Display & User Interface
Color LCD or TFT display
Minimum screen size: 8 inches
Real-time waveform display
Simple menu navigation and parameter adjustment
Display of alarm messages and system status
7. Gas Supply
Compatible with:
Central oxygen and Air supply
Built-in turbine system for operation without compressed air source
8. Power Supply & Battery
Power supply:
220–240 VAC, 50 Hz
Internal rechargeable battery backup:
Minimum 60 minutes operation
Battery charging indicator included
9. Accessories & Consumables
The ventilator shall be supplied with:
Adult breathing circuit
Pediatric breathing circuit
Humidifier chamber (basic)
Reusable flow sensor
Test lung
Oxygen hose and air hose
Power cable
Basic filters and consumables
10. Operating & Storage Conditions
Suitable for Sudan climate conditions:
Operating temperature:
10°C – 40°C
Relative humidity:
20% – 85%
Storage temperature:
5°C – 45°C
11. Standards & Safety Compliance
The ventilator shall comply with:
ISO 13485 or CE certification
IEC 60601-1
IEC 60601-1-2
ISO 80601-2-12 (Critical care ventilators)
12. Training
Supplier shall provide user training covering:
Device operation
Ventilation modes
Alarm management
Maintenance and Troublshooting
Cleaning and basic maintenance
13. Warranty & Maintenance
Minimum 2-years warranty
Preventive and corrective maintenance during warranty period
Spare parts availability for minimum 10 years
14. Documentation
Supplier shall provide:
User manual (English)
Service manual (English)
Spare parts list
Calibration and inspection certificates
3 | C-Arm -Technical Specifications
1.Description of Function
1.1This equipment is used in orthopedic fractures for imaging of bone pathology or fractures on display monitor during operation/ reduction of fractures.
2.Operational Requirements For continuous fluoroscopy, image storage and retrieval the system offered shall be a general fluoroscopy/radiology system
3.System Configuration
3.1X-ray C-Arm Mobile with complete accessories
4.Technical Specifications
4.1X-Ray Generator
X-Ray Generator Microprocessor based, high frequency inverter generator
Generator Output: 3.5kW or more
Fluoroscopic/ Radiographic KV range Lower limit shall not exceed 40 KV Higher limit shall not be less than 110 KV
Fluoroscopic mA range Lower limit shall be ~0.1 mA Upper limit shall be ~9 mA
4.2X-Ray Tube
Rotating anode type
Small focal spot, shall not be more than 0.3 mm
Large focal spot, shall not be more than 0.6 mm
Nominal voltage: 110 KV
Anode heat storage capacity not less than 70 KHU
Inherent filtration should be at least 0.5 mm Al eq
4.3Collimator
Operator controlled automatic collimation
4.4 C-Arm · Focus - I.I. Distance shall be at least 100 cm
· Depth shall be ~ 60cm
· Horizontal travel at least 200mm
· Vertical travel at least 430 mm
· Orbital movement shall be ~120°
· Swivel range shall be ~12°
· Rotation about horizontal axis shall be more than +/-180°
4.5Image Intensifier:
· Triple Mode 9” with direct coupling with camera
· Shall be at least 52 lp/ cm
· Noise reduction, scattered light trap for high contrast dynamics
· CCD camera technology with ABC and AGC control
4.6TV Monitor
· One unit LCD monitor side by side for live and reference image
· Shall be at least 32” with automatic brightness control
4.7Image rotation
· Shall be at least 625 scanning lines at 50 Hz
· Trolley for one display screens and with the alphanumeric keyboard included
· High resolution
4.8 Imaging Modes
· Fluoroscopy mode shall have the following facilities:
· Continuous fluoroscopy with last image hold
· Last image holds with at least two frames image memory
· Continuous fluoroscopy with image acquisition rate: about 20 frame/second.
· Hard disk with image storage capacity of at least 30000 images
· RAM Memory of 256 images
· Mosaic display of 16 images · Zoom (x 2)
· Come with CD/DVD/RW drive
5.System Configuration Accessories, spares and consumables
5.1 Sterilizable textile cover and clips, for the X-ray tube and the Cassette holder for 24 x 30 cm
5.2 All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer.
6. Operating Environment
6.1The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature, Humidity, etc. for Sudan.
6.2 Should work on 220-240V/ 50 Hz AC Single phase fitted with appropriate plugs and sockets.
Standards and Safety Requirements
The unit offered shall be certified to meeting the relevant requirements of TUV, CE mark (MDD), FDA and/ or any equivalent quality and safety standards.
8. User Training
8.1On site operational training till the familiarity of the system and satisfaction of end user shall be provided.
9.Warranty
9.1 Comprehensive warranty for 2 years.
10 Maintenance Service During Warranty Period
10.1Preventive & Corrective Maintenance
11. Installation, Inspections and Commissioning
11.1Supplier must accomplish proper installation and commissioning of the equipment on site
12.Documentation
12.1 User (Operating) manual in English
12.2Service (Technical / Maintenance) manual in English
4 | LED Double Head Ceiling Lamp
1 | Description of Function
1.1 | These lights provide cool, shadow free light and have special technology and filters to provide the same.
2 | Operational Requirements
2.1 | A major operating light, ceiling type with one main & one satellite light units.
3 | System Configuration
3.1 | Operating light ceiling type having dual dome with all standard accessories.
4 | Technical Specifications
I | Main Light, 1 unit.
4.1 | Shall be a ceiling mounted light with flexible arm.
4.2 | Number of light head: min. 1 no.
4.3 | Light head shall not be greater than 760mm diameter.
4.4 | Number of bulb: 1 main bulb, LED with 1 backup bulb with auto-switching or multi-bulbs light head.
4.5 | Light intensity range, shall not be less than 100,000 lux at 1 meter distance from light source
4.6 | Light temperature between 4000 - 4500K.
4.7 | Colour rendering index in between 92 - 94.
4.8 | Shall have 99% heat filtrating.
4.9 | Temperature increase at head shall not be more than 2oC.
4.1 | Temperature increase on operating field shall not be more than 15oC.
4.11 | Radiant heat energy (beam temperature) shall not exceed 25,000 microwatts per square centimetre measured 1m from the light source.
4.12 | The light offered shall have safety designed to prevent patient from burns, especially during the ophthalmic procedures. The light offered shall be certified safe to be used under ophthalmic procedures.
4.13 | Working distance range (focal length): 70 - 180cm.
4.14 | Depth of field with focused light: > 70cm.
4.15 | Luminance field size: 15-25cm diameter, adjustable.
4.16 | Shall have a control to regulate light intensity and to switch on the unit.
4.17 | Shall have an On/Off switch at lamp head.
4.18 | Shall come with continuous dimmer, continuous focus adjustment, continuous field adjustment.
4.19 | Sterilizeable handle to regulate light field size, 3 pcs.
4.2 | Vertical adjustment shall not be less than 115cm.
4.21 | Rotation: 360°.
4.22 | Shall come with Ophthalmic procedures safe light bulbs with a minimum of 1000 hours lifespan.
II | Satellite Light, 1 unit.
4.23 | Shall be ceiling mounted together with the main light with flexible arm.
4.24 | Number of light head: 1 no.
4.25 | Light head shall not be greater than 500mm diameter.
4.26 | Number of bulb: min. 1 main bulb, LED with 1 backup bulb with auto-switching or multi-bulbs light head.
4.27 | Light intensity range, shall not be less than 80,000 lux at 1 meter distance from light source.
4.28 | Light temperature between 4000 - 4500K.
4.29 | Colour rendering index in between 92 - 94. Bidder shall attached certified test certificated showing the compliance of this requirement with TSF.
4.3 | Shall have 99% heat filtrating.
4.31 | Temperature increase at head shall not be more than 2o C.
4.32 | Temperature increase on operating field shall not be more than 15oC.
4.33 | Radiant heat energy (beam temperature) shall not exceed 25,000microwatts per square centimetre measured 1m from the light source.
4.34 | The light offered shall have safety designed to prevent patient from burns, especially during the ophthalmic procedures. The light offered shall be certified safe to be used under ophthalmic procedures.
4.35 | Working distance range (focal length): 70 - 180cm.
4.36 | Depth of field with focused light: > 70cm.
4.37 | Luminance field size: 15-25cm diameter, adjustable.
4.38 | Must have a control to regulate light intensity and to switch on the unit.
4.39 | Shall have an On/Off switch at lamp head.
4.4 | Shall come with continuous dimmer, continuous focus adjustment, continuous field adjustment.
4.41 | Sterilizeable handle to regulate light field size, 3 pcs.
4.42 | Vertical adjustment shall not be less than 115cm.
4.43 | Rotation: 360°.
4.44 | Shall come with Ophthalmic procedures safe light bulbs with a minimum of 1000 hours
4.45 | The main light and satellite light shall comply with interference suppressed VDE 0875 or equivalent.
4.46 | Transformer and operating elements shall be integrated in housing of main light & satellite light.
4.47 | Installation Kit
The followings items shall also be included:
 Ceiling mounting plate/ bracket or equivalent and works and materials to make good the ceiling after installation.
 Wires, conduits and other accessories for connecting the wall control box, the light and others.
 Other materials needed for the installation on the items above.
5 | Accessories, spares and consumables
5.1 | Accessories:
 1 x spare set of fuses.
5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
6 | Operating Environment
6.1 | The system offered shall be designed to store and to operate normally under the conditions include Climate, Power supply , Temperature, Humidity, etc. for Sudan
6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
7 | Standards and Safety Requirements
7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
7.3 | Electrical safety conforms to standards for electrical safety IEC 60601-1 General requirement for Electrical safety of Medical Equipment.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
11 | Installation and Commissioning
11.1 | The bidder must arrange for the equipment to be installed and commissioned by certified or qualified personnel; any prerequisites for installation to be communicated to the purchaser in advance, in detail.
12 | Documentation
12.1 | User (Operating) manual in English.
12.2 | Service (Technical / Maintenance) manual in English.
12.3 | List of important spare parts and accessories with their part numbers and costing.
5 | Operating Table
1.1 | Hydraulic operating tables are simple tables for performing surgical procedures and it works without electrical power.
2 | Operational Requirements
2.1 | OT Table is required for general surgery and shall have X-Ray translucent tops.
3 | System Configuration
3.1 | Operating Table Hydraulic with complete accessories.
4 | Technical Specifications
4.1 | Four section table top with divided foot section.
4.2 | The table shall be mobile on castors with efficient braking system for stability during surgery.
4.3 | Table top must be constructed from a high-pressure laminate to permit x-ray penetration and fluoroscopy.
4.4 | All table positioning, i.e., height, back section, lateral tilt, trendelenburg, and anti-trendelenburg, except foot and head section must be operated hydraulically.
4.5 | Shall have a manual position selector, whose location must be interchangeable between foot and head end.
4.6 | The casings on the frame and centre supporting column must be made of hygienic stainless steel.
4.7 | Mattress must be radio lucent and suitable for fluoroscopy.
4.8 | Dimensions ( approx. +/- 10 % variations):
 Height: 730-1040 mm.
 Side tilt: + 15 degrees.
 Back section adjustment: - 15 degrees to 70 degrees.
 Foot section adjustment: - 90 to 0 degree, detachable.
 Trendelenburg: 25 degree.
 Anti trendelenburg: 25 degree.
 Head section adjustment: -40 to -30 degrees, detachable.
 Maximum width: 555 mm.
 Length: 1950 mm.
5 | Accessories, spares and consumables
5.1 | Accessories:
 Padded arm rest with straps: pair with damps.
 Anesthesia screen with clamps.
 Side supports: pair with clamps.
 Knee crutches: pair with damps.
 X-ray cassette tray.
 Kidney bridge.
 SS bowl with clamps.
 Infusion rod with clamp.
 Legs Support.
Patient weight capacity: not less than 220kg.
5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature,Humidity, etc. for Sudan.
7 | Standards and Safety Requirements
7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA approved product certificate.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
11 | Documentation
11.1 | User (Operating) manual in English.
11.2 | Service (Technical / Maintenance) manual in English.
11.3 | List of important spare parts and accessories with their part numbers and costing.
6 | Defibrillator Machine
Description of Functions
To be used in emergency & critical care departments to meets various resuscitation and monitoring needs.
Technical Specifications
Defibrillator with complete accessories operated with AC power supply and internal battery
Defibrillation function:
It shall be a manual defibrillator for external defibrillation
Able to perform synchronized defibrillation and non-invasive pacing therapy.
Defibrillation energy selection:
External biphasic: 50 - 270J
External Paediatric: 2 - 20J
System shall be user friendly, lightweight and easily transportable.
Waveform shape: biphasic.
The defibrillator paddles shall be easily interchangeable among adult, child, infant and internal paddles. It shall come with at least adult and paediatric paddles.
Can be used for neonatal/paediatric and adult defibrillation.
The unit shall be able to perform defibrillation and monitoring by using disposable electrodes.
Shall have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
Recharge time shall not be held longer than 15 seconds before discharge.
Energy charge & discharge and other selection/control buttons shall be available at the paddle handles.
ECG monitoring function:
Shall have a 3-leads ECG, Lead I, II & III, monitoring capability protected from defibrillation by mean of ECG electrodes and through-the-paddles monitoring
With heart rate display and alarms
With Lead-fault indicator
Shall have an integrated thermal printer/ recorder with paper speed of 25mm/sec
General function:
Shall have LCD that displaying at least dual ECG channel, HR, battery status, shock indicator and various data. Bidder to specify size of LCD screen and the no. of waveforms which can be displayed.
Shall have audio and visual alarms. (Please indicate in the next column type of alarms available)
Shall have HR limit and shockable rhythms alarms
Shall have a rechargeable battery when it is fully charged it shall deliver approximately 40 - 50 discharges or 2 hours of continuous ECG monitoring. Bidder to specify the type of battery used and number of discharge and monitoring hour.
Must have an electrode contact-quality indicator to minimize the risk of ineffective defibrillation.
The unit shall be portable and come with a carrying bag able to keep all required accessories and consumables.
Please indicate the weight in kilogram (KG) of the unit included all accessories and carrying case. It shall be within 8kg
Accessories, spares and consumables
All standard accessories, consumables and parts required to operate the equipment.
· Rechargeable battery, 1 piece on the unit
· Thermal paper x 2 rolls/sets
· Power cord x 1 set
· 3 wire ECG cable x 1 set for ECG monitoring
· Disposable ECG electrodes, 50 pieces
. Paddle for adult and paediatric
· Carry Bag/case x 1 set
Power supply:
Input power supply: 220/240 V AC , 50Hz single phase schuko plug
Environmental factors
Thu unit shall be capable of operating in ambient temperature of 20-50 deg C and relative humidity of 80%
Standards and safety requirements
Should be FDA or CE approved product certificate.
User Training
Must provide user operation and service training
Warranty
Comprehensive warranty for 2 years.
Documentation
User (Operating) manual in English Should provide 2 sets(hardcopy and soft-copy)
Service (Technical / Maintenance) manual in English Should provide 2 sets(hardcopy and soft-copy)
List of important spare parts, accessories and consumables with their part numbers and costing.
Certificate of calibration and inspection from factory.
7 | Major Surgical Instrument Set
1. General Description
1.1 The Major Surgical Instrument Set shall be suitable for general surgical procedures performed in hospitals and operating theatres.
1.2 The set shall include a complete range of instruments required for routine major surgical operations, particularly abdominal and soft tissue surgery.
1.3 All instruments shall be supplied new, unused, and ready for clinical use.
2. Material Requirements
2.1 All instruments shall be manufactured from medical-grade surgical stainless steel.
2.2 The stainless steel shall conform to AISI 410, AISI 420, AISI 440A/440B or equivalent surgical-grade standards.
2.3 Instruments shall be corrosion-resistant and suitable for repeated autoclave sterilization.
2.4 Instruments shall resist pitting, staining, oxidation, and degradation due to repeated cleaning and disinfection.
2.5 Instruments shall have a satin or matte anti-glare finish suitable for operating room conditions.
2.6 All instruments shall be free from defects such as burrs, cracks, sharp edges, and manufacturing imperfections.
2.7 All joints, hinges, ratchets, and locking mechanisms shall operate smoothly and maintain alignment.
2.8 Cutting instruments shall be capable of re-sharpening without loss of performance.
2.9 Each instrument shall be permanently marked with manufacturer identification and catalogue/reference number.
3. Set Components and Quantities
3.1 Forceps and Clamps
3.1.1 Dressing Forceps (14–16 cm) – Qty: 2
3.1.2 Tissue Forceps (14–18 cm) – Qty: 2
3.1.3 Artery Forceps (Straight, 14–20 cm) – Qty: 8
3.1.4 Artery Forceps (Curved, 14–20 cm) – Qty: 8
3.1.5 Kocher Forceps (14–18 cm) – Qty: 4
3.1.6 Allis Tissue Forceps (14–16 cm) – Qty: 4
3.1.7 Babcock Tissue Forceps (16 cm) – Qty: 2
3.1.8 Sponge Holding Forceps (18–24 cm) – Qty: 2
3.1.9 Backhaus Towel Forceps (9–13 cm) – Qty: 4
3.1.10 Rankin-Crile Hemostatic Forceps (16 cm Straight) – Qty: 2
3.1.11 Rankin-Crile Hemostatic Forceps (16 cm Curved) – Qty: 2
3.1.12 Rochester-Pean Hemostatic Forceps (16 cm Curved) – Qty: 2
3.1.13 Rochester-Ochsner Forceps (16 cm Straight) – Qty: 2
3.1.14 Halsted Mosquito Forceps (12.5 cm Straight) – Qty: 6
3.1.15 Halsted Mosquito Forceps (12.5 cm Curved) – Qty: 6
3.1.16 Dissecting Forceps (14.5 cm) – Qty: 2
3.1.17 Dissecting Forceps (20 cm) – Qty: 2
3.1.18 Tissue Forceps (20 cm, 1×2 Teeth) – Qty: 2
3.2 Needle Holders
3.2.1 Needle Holder (16–20 cm) – Qty: 2
3.2.2 TC Mayo-Hegar Needle Holder (18–20 cm) – Qty: 2
3.3 Scissors
3.3.1 Mayo Scissors (Straight, 14–18 cm) – Qty: 2
3.3.2 Mayo Scissors (Curved, 14–18 cm) – Qty: 2
3.3.3 Metzenbaum Scissors (14–18 cm) – Qty: 2
3.3.4 TC Wire Cutting Scissors (12 cm Angled) – Qty: 1
3.4 Scalpel Handles
3.4.1 Scalpel Handle No. 3 – Qty: 2
3.4.2 Scalpel Handle No. 4 – Qty: 2
3.5 Retractors
3.5.1 Richardson Retractor (Multiple sizes) – Qty: 3 sets
3.5.2 Deaver Retractor (300 × 50 mm) – Qty: 2
3.5.3 Volkmann Retractor (4-prong) – Qty: 2
3.5.4 Volkmann Retractor (6-prong) – Qty: 2
3.5.5 Cushing Vein Retractor – Qty: 2
3.5.6 US Army Double-Ended Retractor – Qty: 2 sets
3.5.7 Ribbon Retractor (25 mm) – Qty: 1
3.5.8 Ribbon Retractor (40 mm) – Qty: 1
3.5.9 Balfour Abdominal Retractor – Qty: 1
3.5.10 Langenbeck Retractors (Small) – Qty: 2
3.5.11 Langenbeck Retractors (Medium) – Qty: 2
3.6 Intestinal and Special Instruments
3.6.1 Doyen Intestinal Forceps – Qty: 4
3.6.2 Lane Tissue Forceps – Qty: 2
3.6.3 Grooved Director (14.5 cm) – Qty: 1
3.6.4 Probe with Eye (13 cm) – Qty: 1
3.6.5 Deschamps Ligature Needle (Right) – Qty: 1
3.6.6 Deschamps Ligature Needle (Left) – Qty: 1
3.7 Suction Instruments
3.7.1 Yankauer Suction Tube (28 cm) – Qty: 2
3.7.2 Frazier Suction Tube – Qty: 1
3.8 Trays and Accessories
3.8.1 Kidney Tray – Qty: 2
3.8.2 Perforated Stainless Steel Sterilization Instrument Tray with Lid – Qty: 1
3.9 Equivalent Instruments
3.9.1 Equivalent standard surgical instruments may be accepted provided they meet the same functional and quality requirements.
4. Sterilization Compatibility
4.1 All instruments shall be suitable for steam sterilization (autoclave).
4.2 Instruments shall withstand repeated sterilization cycles without deformation or performance loss.
4.3 Instruments shall be resistant to corrosion, chemicals, disinfectants, and cleaning agents.
5. Storage and Packing
5.1 Instruments shall be supplied in a stainless steel sterilization tray suitable for autoclaving.
5.2 Instruments shall be individually protected and securely packed to prevent damage during transport.
6. Performance Requirements
6.1 Instruments shall ensure smooth operation and precise alignment.
6.2 Ratchets and locking systems shall function securely.
6.3 Cutting instruments shall maintain sharpness under normal clinical use.
6.4 Instruments shall be suitable for routine major surgical procedures.
7. Warranty
7.1 The supplier shall provide a minimum one (1) year warranty against manufacturing defects.
8. Spare Parts and Replacement
8.1 Replacement instruments shall be available upon request.
8.2 The supplier shall ensure availability of spare parts and replacements for a reasonable period.
9. Quality Standards
9.1 ISO 13485 certification or equivalent shall be provided.
9.2 CE Mark and/or FDA approval is preferred.
9.3 Manufacturer shall provide material certification confirming stainless steel grade.
10. Technical Documentation
10.1 Product catalogue and full instrument list shall be provided.
10.2 Material specification including stainless steel grade shall be included.
10.3 Quality certificates (ISO, CE, FDA or equivalent) shall be provided.
10.4 Warranty statement shall be included.
10.5 Country of origin shall be clearly stated.
8 | MINOR SURGICAL SET
This specification defines a Standard Minor Surgical Instrument Set intended for use in primary healthcare facilities, outpatient clinics, emergency rooms, and dressing units.
The set is designed for routine minor surgical procedures including wound management, suturing, dressing, hemostasis, and basic tissue handling.
All instruments shall be reusable, manufactured from medical-grade stainless steel, and suitable for repeated autoclave sterilization.
2. General Technical Requirements
All instruments shall comply with the following minimum requirements:
Medical-grade stainless steel (AISI 420 or equivalent)
Compliance with ISO 7153-1
Autoclavable at 134°C
Satin or matt anti-glare finish
Corrosion-resistant and reusable
Free from manufacturing defects (burrs, cracks, sharp edges)
Ergonomic design suitable for clinical use
Reliable locking mechanisms for forceps and needle holders
Suitable for repeated sterilization cycles without performance degradation
3. Instrument Composition
3.1 Cutting Instruments
Scalpel Handle No. 3 — Qty: 1
Scalpel Handle No. 4 — Qty: 1
Mayo Scissors Straight (14–15 cm) — Qty: 1
Mayo Scissors Curved (14–16 cm) — Qty: 1
Suture Removal Scissors (Spencer type) — Qty: 1
3.2 Hemostatic Instruments
Mosquito Forceps Straight (12–13 cm) — Qty: 2
Mosquito Forceps Curved (12–13 cm) — Qty: 2
Artery Forceps (Pean) Straight (14–16 cm) — Qty: 2
Artery Forceps (Pean) Curved (14–16 cm) — Qty: 2
3.3 Grasping & Tissue Handling Instruments
Dressing Forceps — Qty: 1
Tissue Forceps (Toothed 1×2) — Qty: 1
Dissecting Forceps (Non-toothed) — Qty: 1
Sponge Holding Forceps (Foerster) — Qty: 1
Needle Holder (Mayo-Hegar) (14–16 cm) — Qty: 1
3.4 Exploration & Support Instruments
Surgical Probe / Director — Qty: 1
Backhaus Towel Clamp — Qty: 2
3.5 Dressing Accessories & Containers
Stainless Steel Kidney Dish (500–700 ml) — Qty: 1
Stainless Steel Gallipot (120–200 ml) — Qty: 1
Stainless Steel Instrument Tray (approx. 25 × 20 × 5 cm) — Qty: 1
Autoclavable Sterilization Box with Lid — Qty: 1
4. Packaging Requirements
Instruments supplied in a rigid autoclavable tray or sterilization container
Instruments properly arranged to prevent movement or damage
Clearly labeled contents list included
Batch/lot number for traceability required
Supplied non-sterile but ready for sterilization
Protected against moisture, corrosion, and mechanical damage during transport
5. Quality Standards
The supplier shall provide valid documentation including:
ISO 13485 Certificate (Medical Devices Quality Management System)
ISO 9001 Certificate (preferred)
ISO 7153-1 compliance statement
CE Mark (Medical Device Regulation compliant)
Declaration of Conformity
Certificate of Origin issued by competent authority
All certificates must be issued by recognized certification bodies.
6. Documentation Requirements
The supplier shall submit the following:
Manufacturer catalogue / technical datasheets
Packing list
Warranty certificate (minimum 24 months)
Sterilization compatibility statement (134°C autoclave)
User and maintenance instructions
Delivery note
Inspection and quality control report
7. Inspection and Acceptance Criteria
Final acceptance shall be based on:
Visual inspection of workmanship and finish quality
Verification of quantities and compliance with specification
Functional testing of instruments:
Cutting performance (scissors)
Locking mechanism integrity (forceps/needle holder)
Alignment and grip performance
Verification of packaging integrity
Review of all required documents and certifications
8. Warranty Requirements
Minimum warranty period: 24 months from acceptance
Covers manufacturing defects and material failures
Replacement of defective items at supplier cost
Full after-sales support during warranty period
9. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted, provided they fully comply with the technical requirements and are approved by the procuring authority.
9 | OBSTETRICS & GYNECO (OBS/GYN) SURGICAL SET
1. General Description
This specification defines a Complete Obstetrics & Gynecology Surgical Instrument Set intended for use in:
Maternity wards
Delivery rooms
Operating theatres
Emergency obstetric care units
The set supports procedures including:
Vaginal examination
Assisted delivery procedures
Episiotomy repair
Uterine exploration
Dilatation & curettage (D&C)
Minor gynecological surgical procedures
Emergency obstetric interventions
2. Material Requirements
Medical-grade stainless steel (AISI 420 / AISI 304 or equivalent)
Corrosion resistant and autoclavable at 134°C
Satin / matt anti-glare finish
Smooth surfaces free from burrs or defects
High mechanical strength for repeated surgical use
Durable joints, ratchets, and locking mechanisms
3. INSTRUMENT COMPOSITION (FULL SET)
3.1 Vaginal Examination & Exposure Instruments
Sims Vaginal Speculum, medium (80 × 30 mm blade approx.) — Qty: 1
Cusco Vaginal Speculum, medium size — Qty: 1
Auvard Weighted Vaginal Speculum — Length 25–28 cm — Qty: 1
Breisky Vaginal Retractor — Length 18–20 cm — Qty: 1
3.2 Grasping & Cervical Instruments
Tenaculum Forceps (single hook, 20–22 cm) — Qty: 1
Allis Tissue Forceps (4×5 teeth, 15 cm) — Qty: 2
Ovum Forceps (24–26 cm, fenestrated jaws) — Qty: 2
Sponge Holding Forceps (Foerster, straight/curved, 24–26 cm) — Qty: 2
Ring Forceps (Uterine Dressing Forceps, 24–26 cm) — Qty: 1
3.3 Hemostatic Instruments
Kelly Hemostatic Forceps, straight (14–16 cm) — Qty: 2
Kelly Hemostatic Forceps, curved (14–16 cm) — Qty: 2
Crile Forceps (14–16 cm) — Qty: 2
3.4 Needle Holding & Suturing Instruments
Mayo-Hegar Needle Holder (16–20 cm) — Qty: 2
3.5 Cutting Instruments
Mayo Scissors Curved (14–17 cm) — Qty: 1
Episiotomy Scissors Curved (14–16 cm) — Qty: 1
Umbilical / Cord Scissors (14–16 cm) — Qty: 1
3.6 Uterine Instruments (D&C Set)
Uterine Curette Sharp (sizes 2–10 mm) — Qty: 1 set
Uterine Curette Blunt (sizes 2–10 mm) — Qty: 1 set
Sims Uterine Sound (25–30 cm) — Qty: 1
Hegar Uterine Dilators (1–10 mm graduated set) — Qty: 1 set
3.7 Retractors
Sims Vaginal Retractor — 18–20 cm — Qty: 1
Richardson Retractor — 20–22 cm — Qty: 1
Lateral Vaginal Retractor — 18–20 cm — Qty: 1
3.8 Auxiliary Instruments & Containers
Kidney Dish, stainless steel — Qty: 2
Gallipot, stainless steel (120–200 ml) — Qty: 4
Instrument Tray, stainless steel (40 × 20 × 6 cm) — Qty: 1
Sterilization Box with Lid (autoclavable) — Qty: 1
4. Performance Requirements
Smooth articulation and precise mechanical movement
Secure grip and alignment of jaws
High resistance to repeated sterilization cycles
Reliable cutting performance for scissors
Strong locking mechanisms for forceps
Suitable for obstetric emergencies and surgical procedures
5. Sterilization Requirements
Compatible with autoclave sterilization at 134°C
Resistant to corrosion and chemical disinfectants
No degradation after repeated sterilization cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed layout with identification
Tray includes protective lid and locking system
Full contents list included
Batch/lot traceability required
Supplied non-sterile but ready for sterilization
7. Quality Standards
Supplier shall provide:
ISO 13485 certification
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue with specifications
Full instrument list with dimensions
Material certificates
Warranty certificate
Sterilization compatibility statement
User & maintenance manual
Packing list & inspection report
9. Inspection and Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance
Locking mechanism integrity
Jaw alignment and grip
Verification of full set completeness
Packaging integrity check
Compliance with technical specification
10. Warranty
Minimum warranty: 24 months
Covers manufacturing defects and material failure
Replacement of defective instruments free of charge
After-sales support throughout warranty period
11. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted provided full compliance with this specification and approval of the procuring authority.
10 | Orthopedic Surgical Instrument Set
1. General Description
This specification defines a Complete Orthopedic Trauma Instrument Set for use in operating rooms and orthopedic departments for fracture management, bone fixation assistance, and trauma surgery procedures.
2. Material Requirements
Medical-grade stainless steel (AISI 420 / AISI 410 or equivalent)
Autoclavable at 134°C
Corrosion resistant and reusable
Satin/matt anti-glare finish
High tensile and torsional strength suitable for orthopedic use
Free from defects, cracks, and sharp edges
3. Instrument Composition with Exact Dimensions
3.1 Bone Holding & Reduction Instruments
Bone Holding Forceps (Reposition type, serrated jaws) — 18–24 cm — Qty: 1
Reduction Forceps (Pointed jaw, self-centering type) — 18–22 cm — Qty: 1
Plate Holding Forceps (Dynamic compression type) — 18–20 cm — Qty: 1
Bone Hook (Single sharp hook, ergonomic handle) — 16–18 cm — Qty: 1
3.2 Bone Cutting & Shaping Instruments
Bone Rongeur Forceps (double action, curved jaw) — 16–20 cm — Qty: 1
Bone Cutter / Nibbler (heavy duty) — 18–22 cm — Qty: 1
Osteotomes (Flat, straight, hex handle):
6 mm × 15 cm
10 mm × 15 cm
12 mm × 15 cm
18 mm × 18 cm
Bone Curettes (sharp & blunt, double-ended):
Small: 14 cm
Medium: 16 cm
Large: 18 cm
3.3 Bone Preparation Instruments
Periosteal Elevator (Freer type, double ended) — 14–17 cm — Qty: 1
Bone Awl (cortical perforation, straight) — 14–16 cm — Qty: 1
Depth Gauge (AO standard type) — 14–16 cm — Qty: 1
3.4 Fixation & Implant Support Instruments
Screwdriver Set (AO compatible, hexagonal coupling):
2.5 mm
3.5 mm
4.5 mm
Length: 15–20 cm — Qty: 1 set
Drill Chuck with Key (manual, Jacobs type) — 14–16 cm — Qty: 1
Pin Cutter (K-wire cutter, heavy duty) — 18–20 cm — Qty: 1
3.5 Impact Instruments
Bone Mallet (stainless steel head, wooden or metal handle)
Head diameter: 40–50 mm
Total length: 25–30 cm — Qty: 1
Gigli Saw Handle Set with wire
Handle length: 12–14 cm — Qty: 1 set
3.6 Retracting Instruments
Langenbeck Retractor — 20 × 25 mm blade, length 18 cm — Qty: 1
Hohmann Retractor (various sizes):
Small: 14 cm
Medium: 16 cm
Large: 18 cm
Sharp Retractor — 16–18 cm — Qty: 1 set
3.7 Auxiliary Instruments
Instrument Tray / Sterilization Basket
Size: 40 × 20 × 6 cm (standard orthopedic tray) — Qty: 1
Silicone/foam instrument fixation insert — included
4. Performance Requirements
Precise mechanical alignment of all moving parts
High resistance to bending and torsional stress
Cutting edges capable of clean bone and wire cutting
Secure locking mechanisms for forceps
Stable grip under heavy orthopedic load
Suitable for repeated trauma procedures
Long service life under intensive sterilization cycles
5. Sterilization Requirements
Compatible with autoclave sterilization at 134°C
Resistant to repeated chemical and thermal disinfection
No corrosion or performance degradation after repeated cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed positions with labeling system
Tray includes protective lid and locking system
Each set includes full contents list
Batch/lot number traceability required
Supplied non-sterile but ready for sterilization
Moisture and impact resistant packaging
7. Quality Standards
Supplier shall provide:
ISO 13485 certificate
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (EU MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue with full technical specifications
Complete instrument list with dimensions
Material certification report
Warranty certificate
Sterilization compatibility statement (134°C)
User and maintenance manual
Packing list and QC inspection report
9. Inspection & Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance (rongeur, cutter, scissors)
Locking mechanism strength
Jaw alignment accuracy
Verification of full set completeness
Packaging integrity check
Compliance with all technical specifications
10. Warranty
Minimum warranty: 12 months from acceptance
Covers manufacturing defects and material failure
Free replacement of defective items
Availability of spare/replacement instruments
11. Compliance Statement
Equivalent products from internationally recognized manufacturers may be accepted, provided they fully comply with this specification and are approved by the procuring authority.
11 | ENT Surgical Instrument Set
1. General Description
This specification defines a Complete ENT (Ear, Nose & Throat) Surgical Instrument Set intended for use in:
Hospitals
ENT clinics
Operating rooms
Emergency departments
The set supports diagnostic and surgical ENT procedures including:
Ear examination and procedures
Nasal examination and nasal surgery
Throat and laryngeal procedures
Foreign body removal
Minor ENT surgical interventions
2. Material Requirements
All instruments shall be manufactured from:
Medical-grade stainless steel (AISI 420 / AISI 304 or equivalent)
Corrosion-resistant and suitable for repeated sterilization
Autoclavable at 134°C
Satin/matt anti-glare finish
Smooth surfaces free from defects, burrs, or sharp edges
High durability for repeated clinical use
3. COMPLETE INSTRUMENT COMPOSITION
3.1 Diagnostic ENT Instruments
Head Mirror (adjustable head band type) — Qty: 1
Tuning Fork Set (128 Hz, 256 Hz, 512 Hz) — Qty: 1 set
Tongue Depressor (stainless steel, adult type) — Qty: 2
Auris Otoscope Set (with ear specula) — Qty: 1 set
3.2 Ear Instruments
Ear Specula Set (2 mm – 8 mm assorted sizes) — Qty: 1 set
Ear Curettes (sharp & blunt, assorted sizes) — Qty: 1 set
Jobson Horne Probe with ring curette — Qty: 1
Alligator Ear Forceps (straight, 14–16 cm) — Qty: 1
Aural Forceps (bayonet type, 14–16 cm) — Qty: 1
Ear Syringe (metal, 200–300 ml) — Qty: 1
3.3 Nasal Instruments
Nasal Speculum (Thudicum type, small/medium/large) — Qty: 1 set
Killian Nasal Speculum — Length 14–16 cm — Qty: 1
Tilley Nasal Forceps — Length 15–18 cm — Qty: 1
Hartmann Nasal Dressing Forceps — Length 15–18 cm — Qty: 1
Nasal Septum Elevator — Length 14–16 cm — Qty: 1
Nasal Suction Cannula (Frazier type, 6–12 Fr) — Qty: 1 set
3.4 Throat & Oral Instruments
Boyle-Davis Mouth Gag (adult size) — Qty: 1
Tonsil Holding Forceps — Qty: 1
Magill Forceps (adult, 20–25 cm) — Qty: 1
Laryngeal Mirror Set (assorted sizes) — Qty: 1 set
Tongue Forceps — Qty: 1
3.5 Cutting & General Instruments
ENT Scissors (straight, 14–16 cm) — Qty: 1
Metzenbaum Scissors (curved, 14–18 cm) — Qty: 1
Scalpel Handle No. 3 — Qty: 1
Scalpel Handle No. 4 — Qty: 1
Artery Forceps Straight (14–16 cm) — Qty: 2
Artery Forceps Curved (14–16 cm) — Qty: 2
3.6 Grasping & Holding Instruments
Dressing Forceps — Qty: 1
Tissue Forceps (toothed 1×2) — Qty: 1
Needle Holder (Mayo-Hegar, 14–16 cm) — Qty: 1
Sponge Holding Forceps (Foerster, 20–24 cm) — Qty: 1
3.7 Suction & Auxiliary Instruments
Suction Tube Set (Frazier type, sizes 6–12 Fr) — Qty: 1 set
Kidney Dish (stainless steel) — Qty: 2
Gallipot (120–200 ml) — Qty: 2
ENT Instrument Tray (40 × 20 × 6 cm) — Qty: 1
Sterilization Box with Lid (autoclavable) — Qty: 1
4. Performance Requirements
Smooth mechanical performance of all instruments
Accurate alignment of jaws and tips
High cutting efficiency and durability
Secure grip and ergonomic handling
Suitable for ENT surgical and diagnostic procedures
Resistant to repeated sterilization cycles
5. Sterilization Requirements
Fully compatible with autoclave sterilization at 134°C
Resistant to corrosion and chemical disinfectants
No degradation after repeated sterilization cycles
6. Packaging Requirements
Supplied in rigid stainless steel sterilization tray (40 × 20 × 6 cm)
Instruments arranged in fixed labeled positions
Tray includes lid and locking system
Full contents list included
Batch/lot traceability required
Supplied non-sterile but ready for sterilization
7. Quality Standards
Supplier shall provide:
ISO 13485 certificate
ISO 9001 (preferred)
ISO 7153-1 compliance
CE Mark (MDR compliant)
Declaration of Conformity
Certificate of Origin
8. Documentation Requirements
Manufacturer catalogue and technical datasheets
Full instrument list with specifications
Material certificates
Warranty certificate
Sterilization compatibility statement (134°C)
User & maintenance manual
Packing list and inspection report
9. Inspection and Acceptance Criteria
Visual inspection (finish, defects, corrosion)
Functional testing:
Cutting performance
Jaw alignment
Locking mechanisms
Verification of full set completeness
Packaging integrity check
Compliance with technical specifications
10. Warranty
Minimum warranty: 24 months
Covers manufacturing defects and material failure
Free replacement of defective instruments
Full after-sales support during warranty period
11. Compliance Statement
Equivalent instruments from internationally recognized manufacturers may be accepted provided they fully comply with this specification and are approved by the procuring authority.
12 | Autoclave (50 L)
1 | Description of Function
1.1 | Autoclaves are required for sterilizing an object in high temperature and high pressure steam.
2 | Operational Requirements
2.1 | Microprocessor based electrically heated vertical steam sterilizer
3 | System Configuration
3.1 | Microprocessor based Autoclave (Vertical Model) with complete accessories.
4 | Technical Specifications
capacity: approx. 50 L.
Vertical type.
Stainless steel.
Digital controlled temperature and pressure system.
Steam sterilization, up to 135° C
Digital temperature and pressure gauges.
Safety devices : over heat ( low water cut-off switch , safety valve and release valve)
Automatic controlled sterilization cycle.
2 modes sterilization (121°c - 134°c )
Exhaust system.
Stainless steel basket.
Double wall case.
5 | Accessories, spares and consumables
Spare heating element- 2 set
All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Climate ,Temperature , Power supply , Humidity, etc. for Sudan.
Power supply: 220-240V/ 50 Hz AC Single phase with appropriate plug .The power cable must be minimum 3 Meter
7 | Standards and Safety Requirements
7.1 | Must submit ISO 9001 or ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE or USFDA approved product certificate.
Shall meet IEC 61010-2-040 Safety requirements for electrical equipment for measurement, control and laboratory use - Part 2-040: Particular requirements for sterilizers and washer-disinfectors used to treat medical materials.
8 | Warranty
8.1 | Comprehensive warranty for 2 years.
9 | Maintenance Service During Warranty Period
9.1 | During warranty period supplier must ensure corrective/breakdown maintenance whenever required.
10 | Installation and Commissioning
10.1 | The supplier must accomplish proper commissioning of the item onsite.
11 | Documentation
11.1 | User and/or service manual shall be supplied in English.
11.2 | User (Operating) manual in English
11.3 | Service (Technical / Maintenance) manual in English
11.4 | List of important spare parts and accessories with their part numbers and costing.
11.5 | Certificate of calibration and inspection from factory.
13 | Infusion Pump
1 | Description of Function
1.1 | A microprocessor controlled infusion pump unit is needed to include but not limited to the following features:
1.2 | Flat hygienic touch screen.
1.3 | Syringe loading sensor – with KVO (keep vein open)
1.4 | Self calibrated, self diagnosis capability
1.5 | Volume range from 1 –999 ml/hr or better in 1 ml increment
1.6 | High accuracy rate< /- 2%
1.7 | Audio visual indicators
1.8 | Multi types A/V alarms to include occlusion, door open, low battery, empty, etc…
1.9 | Open system using standard IV lines
1.1 | Air in line/ fluid detector
1.11 | Built in rechargeable battery, at least two hours operation
1.12 | Clamp pole
2 | Operating Environment
2.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
2.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
7 | Standards and Safety Requirements
7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required.
11 | Documentation
11.1 | User (Operating) manual in English.
11.2 | Service (Technical / Maintenance) manual in English.
11.3 | List of important spare parts and accessories with their part numbers and costing.
14 | Syringe Pump
1 | Description of Function
1.1 | The Syringe Infusion Pump provides uniform flow of fluid by precisely driving the plunger of a syringe down its barrel. It provides accurate and continuous flow rate for precise delivery of I.V. medication in critical medical care.
2 | Operational Requirements
2.1 | The syringe pump must be programmable, user friendly, safe to use and must have battery backup and comprehensive alarm system. This must be able to integrate in the HIS.
3 | System Configuration
3.1 | Syringe infusion pump with battery backup alarm and with complete accessories.
4 | Technical Specifications
4.1 | Flow rate programmable from 0.1 to 200 ml/hr. or more in steps of 0.1 ml/hr. with user selectable flow set rate option. SAVE last infusion rate even when the AC power is switched OFF.
4.2 | Bolus rate must be programmable to 400 – 500 ml/hr. or more with infused volume display. Reminder audio after every 0.5 ml delivered bolus. SAVE last Bolus rate even when the AC power is switched OFF.
4.3 | Display of Drug Name with a provision of memorizing 10~15 names by the operator
4.4 | Keep Vein Open (KVO) must be available 1.0 ml/hr. or set rate if lower than 1.0 ml. User must have choice to disable KVO whenever desired.
4.5 | Selectable Occlusion pressure trigger levels selectable from 300/500/900 mmHg
4.6 | Must Work on commonly available 20, 50/60 ml Syringes with accuracy of minimum of +/-2% or better.
4.7 | Automatic detection of syringe size & proper fixing. Must provide alarm for wrong loading of syringe such as flanges out of slot; disengaged plunger, unsecured barrel etc.
4.8 | Anti-bolus system to reduce pressure on sudden release of occlusion
4.9 | Must have comprehensive alarm package including: Occlusion limit exceed alarm ,Near end of infusion pre-alarm & alarm, Volume limit pre-alarm & alarm, KVO rate flow, Low battery pre-alarm and alarm, AC power failure, Drive disengaged and preventive maintenance.
4.10 | Rechargeable Battery having at least 5~6 hour backup for about 5ml/hr. flow rate with 50ml syringes. Larger battery life and indication of residual life will be preferred.
5 | Accessories, spares and consumables
5.1 | Accessories:
 Mounting device/ Docking Station for two or four pumps as per requirement so as to enable to power up to 2-4 pumps with one power cord when mounted on IV pole. -01 pc.
5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer. Bidders must specify the quantity of every item included in their offer (including items not specified above).
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
7 | Standards and Safety Requirements
7.1 | Must submit ISO13485:2003/AC:2007 for Medical Devices AND
7.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
7.3 | Certified for meting IEC60601-2-24:Particular requirements for the safety of infusion pumps and controllers
7.4 | Must meet IEC 529 Level 3 (IP3X) (spraying water) for enclosure protection, water ingress.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
11 | Documentation
11.1 | User (Operating) manual in English.
11.2 | Service (Technical / Maintenance) manual in English
11.3 | List of important spare parts and accessories with their part number and costing.
15 | Manual Patient Bed One Movement With Mattress
1 | Description of Function
1.1 | A hospital bed is a bed specially designed for hospitalized patients in need of patient ease. These beds have special features both for the comfort and well-being of the patient and for the convenience of hospital staff.
1.2 | Mattress is to provide a comfortable platform to rest or sleep upon the bed.
2 | Operational Requirements
2.1 | The patient bed shall be made of solid steel construction with anti-corrosive and antirust treated baked hard epoxy powder coating
3 | System Configuration
3.1 | Hospital Bed epoxy powder coated
4 | Technical Specifications
4.1 | Bed base shall be anti-corrosive and antirust treated epoxy powder coated welded steel bar or epoxy powder coated 18G perforated sheet top to improve ventilation.
4.2 | The patient bed shall be fixed height with 2 sections where the backrest section could be elevated by mechanical hand crank located at the foot-end of the bed.
4.3 | Shall have 4 IV rod receptacles and mosquito net pole receptacles at the 4 corners
4.4 | It shall come with one dual hook anti-corrosive and antirust treated epoxy powder coated or 304 grade stainless steel IV rod.
4.5 | Shall have provisions to fix urinary bag on both sides.
4.6 | It shall mobile on 4 x 200mm robust swivelling castors with non-marking grey tyres and with at least 2 diagonal castors shall have locking/brake mechanism.
4.7 | Bedhead and foot-end panel (head and foot bows) shall be made of either between 4-6 anti-corrosive and antirust treated epoxy powder coated vertical tubular tube of not less than 30mm in diameter or epoxy powder coated steel panel
4.8 | Both bedhead and foot-end panel shall be detachable.
4.9 | The height of the bedhead panel: not less than 1060mm from floor.
4.1 | The height of the foot-end panel: not less than 820mm from floor.
4.11 | Overall approximate dimension: not less than 1980mm length, 910mm width, 600mm height
4.12 | The mattress negotiable with hospital beds having an overall approximate dimension of not less than 1980mm length, 910mm width.
4.13 | It shall be fire retardant antibacterial treated high density approx. 40kg/m3 PU foam mattress.
4.14 | The mattress shall have thickness of at least 100mm.
4.15 | Mattress with two sections shall be designed to bend with the positioning of the bed when the backrest and foot section of the bed are adjusted.
4.16 | The weight capacity of the mattress shall be more than 100kg.
4.17 | Mattress Cover:
The mattress shall come with a zipped fire retardant antibacterial, antistatic, acid resistance, waterproof and washable vinyl or vinylized nylon cover. It shall be designed to provide ventilating airflow over the patient’s skin. The zip shall be a heavy-duty/large toothed synthetic zipper to enable inspection of the inner foam and totally covered by a flap extending over the zip to prevent ingress of fluids into the actual mattress and allow for replacement.
5 | System Configuration Accessories, spares and consumables
5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer. Bidders shall specify, in a separate Excel worksheet, the quantity and details of any items included in this offer which have not been specified in this Technical Specifications Form.
6 | Operating Environment
6.1 | The system offered shall be designed to store and to operate normally under the conditions. The conditions include Climate, Temperature, Humidity, etc.
7 | Warranty
7.1 | Warranty for 2 years.
16 | Delivery Table
1 | Description of Function
1.1 | Delivery bed is used for Baby Delivery and must incorporate ideal blend of the patient's individual requirements on comfort and the professional needs of the delivery team, focusing on the aesthetic and functional design of the entire product.
2 | Operational Requirements
2.1 | Manually operated delivery bed.
3 | System Configuration
3.1 | Delivery Bed with complete attachments and accessories.
4 | Technical Specifications
4.1 | It must have manual adjustments for height and back positions.
4.2 | It must have collapsible side rails.
4.3 | It must have three sectional mattresses and seat section must have large perennial cut.
4.4 | It must have head board which can be detached.
4.5 | Must have wheels provided with locking system.
4.6 | Must have retractable foot section so as to convert bed into table.
4.7 | Must have infusion rods which have adjustable heights, quick release and attaches to all corners of bed.
4.8 | Must have adjustable leg rests.
4.9 | Must have push grip handles.
4.10 | Must have sliding stainless steel bowl at perennial part of table.
4.11 | It must have catheter bag holder which can be attached on either side of bed.
4.12 | It must be able to give trendelenburg, reverse trendelenburg and 70 degree sitting position.
4.13 | It must have adjustable foot supports.
4.14 | It must be easy to maintain clean and sterilize (especially blood stains).
4.15 | Frame must be (washable) stainless steel.
4.16 | Dimensions (approx.):
 Length: 7 feet
 Width: 3 feet
4.16 | Capacity load of 180 kg or more
5 | Accessories, spares and consumables
5.1 | All standard attachments and accessories: 01 set
5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature , Humidity , etc. for Sudan.
7 | Warranty
7.1 | Comprehensive warranty for 2 years after acceptance.
8 | Maintenance Service During Warranty Period
8.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
17 | Examination Couch
1 | Description of Function
1.1 | Examination couch for use of health checkup and treatment of patients.
2 | Operational Requirements
2.1 | An examination couch with upholstered top in two pieces. Adjustable headrest on gas spring.
3 | System Configuration
3.1 | Examination couch with mattress.
4 | Technical Specifications
4.1 | The examination couch shall be made of a solid steel sheet and plate construction with anti-corrosive and antirust treated epoxy powder coating with upholstered top.
4.2 | All 4 legs of the bed shall be capped with heavy duty rubber footings.
4.3 | Overall size of the table must not be less than 1890mm L x 600mm W x 825mm H
4.4 | Strong Mild steel tubular construction epoxy powder coated treated. The top base of machine pressed double bent Mild steel sheet epoxy powder coated treated finish
4.5 | Gas spring assisted adjustable backrest of approx. size 450mm L x 310mm H with upholstered top.
4.7 | Swinging tray must be attached near headrest for BP apparatus and/or other health checkup minor equipment.
4.8 | The mattress shall be foldable and shall be designed to bend with the positioning of the bed when the backrest of the bed is adjusted.
4.9 | Bidder shall indicate the weight capacity and the total weight of the mattress in kilogram (kg)
4.1 | The mattress shall have mid-firmness, with foam density of approximately 0.55kg/ cubic foot, to avoid that the patient would sink down into foam with antibacterial, antistatic, acid resistance, waterproof and washable vinyl or vinylized nylon cover.
4.11 | The joints must be smooth and neat finish.
5 | Accessories, spares and consumables
5.1 | All standard accessories/consumables/parts required for the proper operation of the above item shall be included in the offer, the quantity and details of any items included in this offer which have not been specified in this Technical Specifications Form.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Climate, Temperature,Humidity, etc. for Sudan.
7 | Warranty
7.1 | Comprehensive warranty for 2 years after acceptance.
8 | Maintenance Service During Warranty Period
8.1 | During the warranty period supplier must ensure corrective/breakdown maintenance whenever required.
16 | Portable Ultrasound Machine
1 | Description of Function
1.1 | A general purpose fully digital B & W Ultrasound imaging system.
2 | Operational Requirements
2.1 | It shall operate on mains AC supply.
3 | System Configuration
3.1 | System shall come with main unit, 1 probe, 1 unit of black and white thermal printer and Ultrasound gel warmer 1 unit.
4 | Technical Specifications
4.1 | Latest technology and all digital beam former general purpose standalone ultrasound machine with integrated light weight mobile cart.
4.2 | Main applications: OB/GYN, abdominal, peripheral vessels and small parts.
4.3 | The system shall have at least 12” or higher flat LCD monitor with tilt & swivel facilities.
4.4 | Shall have B-mode, M-mode, B/M mode, 2B mode & 2D mode.
4.5 | The system must have at least Two active probe ports for easy use and convenient operation.
4.6 | 256 Grey shades for sharp contrast resolutions.
4.7 | Controls for depth, gain compensation, body markers with transducer position.
4.8 | Shall have real time continuous dynamic focus.
4.9 | Shall have facility for image zoom, freeze, text annotation.
4.1 | The system shall have extensive calculation software package for Ob/Gyn and general imaging.
4.11 | The system must have provision for measurement and calculation of distance, area, volume, heart rate and circumference on the image.
4.12 | The system shall have Tissue Harmonic Imaging.
4.13 | Near and far gain adjustable.
4.14 | Contrast, adjustable.
4.15 | Focus: auto adjustable.
4.16 | Shall have an alpha-numeric keyboard with easy access scans controls and track ball and status display.
4.17 | Cine memory of 250 frames for cine loop playback.
4.18 | Frame rate: not less than 50fps.
4.19 | Display depth: minimum 28-30cm.
4.2 | Dynamic range, selectable up to approximately 165dB.
4.21 | Image storage: Minimum 200 patient’s images on main unit.
4.22 | Shall have facility for inbuilt CD writer.
4.23 | System shall be DICOM
4.24 | Facility for future upgradeability.
4.25 | Probe: 2 to 5 MHz convex probe for Obs. /Gyn. and abdominal application is to be supplied.
5 | Accessories, spares and consumables
5.1 | Accessories:
 Black and white thermal printer with 50 rolls of high density recording paper: 01 no.
 DVD/CD Recorder with DICOM media transfer.
 Ultrasound gel warmer: 01 unit.
5.2 | All standard accessories, consumables and parts required to operate the equipment, including all standard tools and cleaning and lubrication materials, to be included in the offer.
6 | Operating Environment
6.1 | The product offered shall be designed to be stored and to operate normally under Power Supply, Climate, Temperature,Humidity, etc. for Sudan.
6.2 | Power supply: 220 – 240 VAC, 50Hz fitted with appropriate plug. The power cable must be at least 3 metre in length.
6.3 | Shall provide UPS of suitable rating with voltage regulation and spike protection for 30 minutes back-up.
7 | Standards and Safety Requirements
7.1 | Must submit ISO 13485:2003/AC: 2007 AND
7.2 | CE (93/42 EEC Directives) or USFDA or TUV approved product certificate.
7.3 | Electrical safety conforms to standards for Electrical Safety IEC 60601-2-37 Medical electrical equipment – Part 2-37: Particular requirements for the basic safety and essential performance of ultrasonic medical diagnostic and monitoring equipment.
8 | User Training
8.1 | Must provide user training (including how to use and maintain the equipment).
9 | Warranty
9.1 | Comprehensive warranty for 2 years after acceptance.
10 | Maintenance Service During Warranty Period
10.1 | During the warranty period supplier must ensure preventive maintenance and corrective/breakdown maintenance whenever required.
11 | Documentation
11.1 | User (Operating) manual in English.
11.2 | Service (Technical / Maintenance) manual in English.
11.3 | List of important spare parts and accessories with their part number and costing.