## Sheet1
1 | Video Gastroscope
No. | Item Specifications | Fill your Specification
1.1 | Field of view 140°
1.2 | Observation range 2 - 100 mm
1.3 | Bending capability Up210°/Down90° Right 100°/Left 100°
1.4 | Distal end ø 9.2 mm
1.5 | Insertion tube ø
1.6 | Working channel ø 2.8 mm
1.7 | Working length 1100 mm
1.8 | Total length 1400 mm
1.9 | Must be compatible with the video system specified
2 | Video Duodenoscope
No. | Item Specifications | Fill your Specification
1.1 | Direction of view must be zero degree.
1.2 | Field of View: 100°.
1.3 | Observation Range: 4 – 60 mm
1.4 | Distal end diameter approx 13.1 mm
1.5 | Insertion tube diameter of less than 11.5 mm.
1.6 | Angulations of tip up 130° and down 90° with right and left movement of at least 110° /90° .
1.7 | Inner diameter of instrument channel approx 4,2 mm
1.8 | Working length approx of 1250mm
1.9 | Total length approx of 1550mm
1.10 | Must be compatible with the video system specified
3 | Video Colonoscope
No. | Item Specifications | Fill your Specification
1.1 | Long 1.6 m
1.2 | Direction of view must be zero degree.
1.3 | Field of View: 140°.
1.4 | Observation Range: 2 – 100 mm
1.5 | Distal end diameter of not more than 12.8 mm
1.6 | Insertion tube diameter of less than 12,8 mm.
1.7 | Angulations of tip up 180° and down 180° with right and left movement of at least 160° /160° .
1.8 | Instrument channel minimum 3.8mm
1.9 | Working Length approx: 1700mm
1.10 | Total length approx: 2000mm
1.11 | Must be compatible with the video system specifie
1 | Option 1
1) High-Definition Endoscopy Video Processor (separate light source)
No. | Item Specifications | Fill your Specification
1 | Imaging & Resolution
1.1 | Native output: Full HD 1080p (1920×1080) or higher; progressive scan.
1.2 | Input compatibility: HD scopes and legacy SD scopes via appropriate adapters.
1.3 | Frame rate: ≥50/60 fps at 1080p.
1.4 | Signal processing: Real-time noise reduction, dynamic contrast enhancement, edge enhancement, color balance, and automatic white balance.
1.5 | Digital zoom: ≥1.5× (stepwise or continuous) without frame drop.
1.6 | Latency: ≤50 ms end-to-end at 1080p.
2 | Image Enhancement / Virtual Chromoendoscopy (brand-neutral wording)
2.1 | At least one integrated spectral/virtual enhancement mode for mucosal/vascular detail (e.g., narrow-band/spectral filtering or post-processing).
2.2 | Quick toggle from front panel and scope button; on-screen indicator of mode.
3 | Connectivity & Outputs
3.1 | Digital video: ≥2× HDMI or DVI-D outputs (1080p).
3.2 | SDI or DisplayPort (optional).
3.3 | USB 3.0 ports (≥2) for image/video capture to external media.
3.4 | Network: Gigabit Ethernet for PACS/VNA/DICOM (Store-SCU) or HL7 gateway; optional Wi-Fi.
3.5 | DICOM stills/video export with patient demographic overlay.
3.6 | Remote service port (LAN/USB) for diagnostics/logs.
4 | Recording & Data Management
4.1 | Still image: JPEG/PNG, full resolution.
4.2 | Video: H.264/H.265, 1080p, selectable bitrate; simultaneous live output & recording.
4.3 | Patient data entry via on-screen keyboard, barcode scanner, or HIS/RIS worklist (optional).
4.4 | Storage: Internal ≥256 GB (optional), plus external USB.
4.5 | Metadata: Date/time, scope ID, operator ID, enhancement mode, and procedure type embedded.
5 | User Interface & Controls
5.1 | Front panel: Backlit, wipeable keys; rotary dial for gain/contrast.
5.2 | Programmable scope buttons (≥2) for capture, enhancement toggle, freeze.
5.3 | On-screen display: Battery/PSU status (if applicable), light level, enhancement mode, recording status.
6 | Compatibility
6.1 | Must connect to the proposed light source via standard light-guide port and to manufacturer’s endoscopes (state supported connector types: e.g., “O”, “F”, “P” types).
6.2 | Backward compatibility with existing facility scopes (list models in tender addendum).
7 | Safety & Standards
7.1 | Electrical safety: IEC 60601-1; EMC: IEC 60601-1-2; Medical software lifecycle: IEC 62304 (declared).
7.2 | Ingress protection for front panel: ≥IP21.
7.3 | Medical CE mark and/or FDA 510(k)/TGA/other local market authorization (specify jurisdiction).
8 | Physical & Power
8.1 | Power: 100–240 V AC, 50/60 Hz; power consumption ≤150 W (processor alone).
8.2 | Dimensions: ≤45 cm (W) × 12 cm (H) × 40 cm (D) (or equivalent rack-mountable 3U).
8.3 | Noise: ≤30 dB(A).
9 | Accessories (Included)
9.1 | HDMI cable (certified 1080p), DVI/HDMI adapters as needed, 2× USB 3.0 memory sticks (≥64 GB), barcode scanner (optional), footswitch (optional).
10 | Service & Warranty
10.1 | Warranty: ≥24 months, extendable to 60.
10.2 | Preventive maintenance: ≥2 visits/year or per OEM schedule.
10.3 | Local service response: ≤72 hours; loan unit provision (optional).
10.4 | User training: Initial on-site training for physicians/nurses/biomed, with certificates; training materials (videos/manuals).
2) Endoscopic Light Source (separate unit)
No. | Item Specifications | Fill your Specification
1 | Illumination Technology (select one or allow both)
1.1 | LED: High-power white LED array; expected lamp life ≥20,000 hours.
1.2 | Xenon: 300 W short-arc xenon lamp; lamp life ≥500 hours with hot-swap or quick replacement.
2 | Optical Performance
2.1 | Luminous output at light-guide port:
2.2 | LED: ≥700–1,000 lumens equivalent (state OEM test method).
2.3 | Xenon: ≥1,000 lumens.
2.4 | Color temperature: 5,600–6,500 K (daylight range).
2.5 | Uniformity: ≥85% across field; minimal hotspotting.
2.6 | Automatic brightness control (ABC) linked to processor video feedback; manual override with ≥10 steps.
2.7 | Fine iris/aperture control for close-up work.
3 | Specialty Illumination / Filters (if enhancement mode requires)
3.1 | Built-in optical filters or electronic control compatible with the processor’s enhancement mode(s) (e.g., spectral narrow-band).
3.2 | Quick-select button for enhancement illumination; status indicator.
4 | Air/Water Integration (if applicable to model)
4.1 | Integrated medical-grade air pump for insufflation and water bottle feed, flow adjustable; or compatibility with external CO₂ insufflator (preferred for GI).
4.2 | Filtered air path (HEPA or equivalent) with replaceable filters; water path materials biocompatible.
5 | Safety & Protections
5.1 | Over-temperature monitoring with auto dim/shutdown and audible alarm.
5.2 | Lamp life counter (hours used & remaining); replacement reminder.
5.3 | Interlock to prevent light port exposure (shutter closes when scope disconnected).
5.4 | Compliance: IEC 60601-1, 60601-1-2, photobiological safety IEC 62471.
6 | Connectors & Compatibility
6.1 | Light-guide port to match the facility’s scopes (declare connector type(s)); include adapters if mixed fleet.
6.2 | Secure bayonet or threaded coupling; strain relief for light cable.
6.3 | Compatible with the specified processor’s ABC.
7 | Physical & Power
7.1 | Power: 100–240 V AC, 50/60 Hz.
7.2 | Typical consumption:
7.3 | LED: ≤120 W at max.
7.4 | Xenon: ≤400 W at max.
7.5 | Cooling: Quiet forced-air; noise ≤40 dB(A).
7.6 | Dimensions: ≤45 cm (W) × 12 cm (H) × 40 cm (D).
8 | Accessories (Included)
8.1 | 1× standard light-guide cable matching scope connector, length 2.5–3.0 m, autoclavable or low-temperature sterilizable as per OEM.
8.2 | 1× spare lamp module (for Xenon) or commitment for rapid lamp supply; filter set if required.
9 | Warranty & Service
9.1 | Same terms as (Section 10), coverage per OEM policy.
1 | Option 2
Combined HD Endoscopy Video Processor & Light Source (“Combo Unit”) — General Specifications
No. | Item Specifications | Fill your Specification
1 | Imaging & Resolution
1.1 | Native output: Full HD 1080p (1920×1080) or higher (progressive scan).
1.2 | Frame rate: ≥50/60 fps (progressive).
1.3 | Image processing: Real-time automatic gain, contrast, gamma, and white balance.
1.4 | Enhancement modes: Integrated electronic image enhancement / virtual chromoendoscopy (narrow-band / spectral / tone enhancement) for mucosal and vascular detail.
1.5 | Noise reduction: Adaptive 2D/3D noise reduction and edge enhancement.
1.6 | Zoom: ≥1.5× optical/digital.
1.7 | Freeze & snapshot: Instant capture via scope or panel button.
2 | Light Source (Built-in)
2.1 | Type: High-intensity LED array (preferred for long life and low heat) or optionally Xenon 300 W short-arc in larger models.
2.2 | Luminous output: ≥800 lumens (LED) / ≥1,000 lumens (Xenon).
2.3 | Color temperature: 5,600–6,500 K (daylight equivalent).
2.4 | Lamp life:
2.5 | LED ≥20,000 hours (maintenance-free).
2.6 | Xenon ≥500 hours, quick replacement cartridge.
2.7 | Brightness control: Automatic (ABC) via video feedback, manual override knob (≥10 steps).
2.8 | Enhancement lighting: Compatible with spectral filter or digital enhancement mode.
2.9 | Safety: Automatic light shutter when scope is disconnected.
3 | Connectivity & Outputs
3.1 | Video output: ≥2× HDMI or DVI-D (1080p).
3.2 | Optional: 3G/12G-SDI or DisplayPort for medical monitor connection.
3.3 | USB ports: ≥2 (USB 3.0) for image/video recording.
3.4 | Network: LAN (Ethernet, Gigabit) for DICOM image export to PACS.
3.5 | Audio output: Optional line-out for procedure room recording.
3.6 | Footswitch: Optional (start/stop capture).
4 | Recording & Storage
4 | Still images: JPEG or PNG, full HD.
4.1 | Videos: H.264 or H.265 compression, 1080p, 30/60 fps.
4.2 | Internal memory: ≥256 GB SSD (optional).
4.3 | External storage: USB flash drive or external HDD.
4.4 | Patient data overlay: ID, procedure, time/date, operator name.
5 | Display Compatibility
5.1 | Output suitable for 21–32″ medical-grade monitors (1080p).
5.2 | Optional 4K upscaling (for new 4K monitors).
5.3 | Adjustable aspect ratio 16:9 / 4:3.
6 | User Interface
6.1 | Front panel: Touchscreen or membrane keys (waterproof).
6.2 | Language menu: English + multilingual interface.
6.3 | Programmable scope buttons: ≥2 functions assignable (capture, enhancement toggle).
6.4 | Status indicators: Lamp on/off, brightness level, recording status, error alarms.
7 | Compatibility
7.1 | Compatible with endoscopes of the same manufacturer (declare model type).
7.2 | Adapters available for other standard connectors (O, F, P types).
7.3 | Optional compatibility with rigid scopes or ENT scopes via coupler.
8 | Safety & Standards
8.1 | Medical electrical safety: IEC 60601-1.
8.2 | EMC compliance: IEC 60601-1-2.
8.3 | Light safety: IEC 62471 (photobiological safety).
8.4 | CE mark / FDA 510(k) / TGA or equivalent regulatory approval.
8.5 | Built-in overheat protection and light intensity limiter.
9 | Physical Characteristics
9.1 | Power supply: 100–240 V AC, 50/60 Hz (universal).
9.2 | Power consumption: ≤250 W (LED type) / ≤450 W (Xenon type).
9.3 | Dimensions: ≤45 cm (W) × 15 cm (H) × 40 cm (D).
9.4 | Weight: ≤10 kg.
9.5 | Cooling: Quiet forced-air cooling (<40 dB).
10 | Accessories
10.1 | 1× Autoclavable light guide cable (length 2.5–3.0 m).
10.2 | 1× Power cable and HDMI cable.
10.3 | 1× User manual and quick-start guide.
10.4 | 1× Spare lamp (for Xenon model only).
10.5 | Optional: footswitch, video capture software license, DICOM configuration.
11 | Service & Warranty
11.1 | Warranty: Minimum 2 years for system
11.2 | Preventive maintenance visits per manufacturer schedule.
11.3 | Local service support with response ≤72 h.
11.4 | Training: On-site demonstration and user certification for medical/nursing staff.