## Sheet1
RFQ 12-2025- Specs
1 |  | Digital Radiography (X-Ray Machine)
| No. | Item Specifications | Fill your Specification
| 1 | Description of Function
| 1.1 | Digital X-ray machine for routine X-ray examinations at healthcare facilities.
| 2 | Operational Requirements
| 2.1 | It shall be suitable to be used for adult and pediatric patients in general Radiography examination.
| 3 | System Configuration
| 3.1 | X-ray Generator,1 unit
| 3.2 | X-Ray tube & tube support system, 1 unit
| 3.3 | Radiographic patient table, fixed height1 unit
| 3.4 | Floor mounted bucky stand, 1 unit
| 5 | Technical Specifications
| 5.1 | Should be floor to floor mounted
| 5.2 | X-ray Generator:
| 5.2.1 | Microprocessor based, high frequency inverter generator, the generator shall have at least 100kHz.
| 5.2.2 | Generator Output: not less than 30 Kw
| 5.2.3 | Radiographic voltage: 40 kV to 140kV, in 1Kv step.
| 5.2.4 | Radiographic current approx., 10 to 500 mA
| 5.2.5 | mAs range: 0.5 mAs to 650 mAs.
| 5.2.6 | Anatomical Programmable Radiographic mode shall be available.
| 5.2.7 | Manual & automatic exposure control and automatic brightness control shall be available.
| 5.2.8 | Shall come with overload protection device.
| 5.3 | X-Ray Tube:
| 5.3.1 | Maximum tube output shall match with the generator output
| 5.3.2 | X-ray tube rotating: +/-90°.
| 5.3.3 | Anode rotating speed: 2700 rpm or more.
| 5.3.4 | Large focus not more than 1.2 mm.
| 5.3.5 | Small focus not more than 0.6 mm.
| 5.3.6 | Maximum tube voltage 140 KV.
| 5.3.7 | Max housing heat dissipation rate: 250HU/s
| 5.3.8 | Anode heat storage capacity: min 140 kHU (100kJ);
| 5.3.9 | Max. heat storage capacity of the tube housing: 9,000kJ (1,250kHU).
| 5.3.10 | Cooling method passive or forced air and/or oil cooling.
| 5.3.11 | Filtration: min 2.7 mm Al equivalent.
| 5.4 | Floor Mounted Tube Stand:
| 5.4.1 | Horizontal travel range: 138 cm to 220 cm, movement arrested by electromagnetic brakes;
| 5.4.2 | Vertical travel range: 150 cm to 160 cm movement arrested by electromagnetic brakes;
| 5.4.3 | Central beam height: min 35 cm to min 185 cm;
| 5.4.4 | Rotation of tube around vertical axis: ± 90°; stop positions: 0°, ± 90°;
| 5.4.5 | X-ray tube rotation: ± 120°; stop positions: 0°, ± 90°.
| 5.5 | Flat Panel Detector:
| 5.5.1 | Dimensions (active field): min 34 cm x max 43 cm;
| 5.5.2 | Detector thickness: 16 mm
| 5.5.3 | Pixel size: 139 µm;
| 5.5.4 | Spatial resolution (Nyquist frequency): 3.6 lp/mm.
| 5.6 | Collimator:
| 5.6.1 | Inherent filtration: min 1 mm Al at 70 kV;
| 5.6.2 | Full-field light localizer: 150 W halogen light, timer;
| 5.6.3 | Built-in light switch should be provided.
| 5.6.4 | Copper prefilter: Without filter, 0.1 mm, 0.2 mm, 0.3 mm Cu;
| 5.6.5 | Rotation: Up to a maximum of ± 45°;
| 5.6.6 | Collimation control: Manual.
| 5.7 | Floor Mounted Bucky Stand:
| 5.7.1 | Vertical movement range: min 150 cm;
| 5.7.2 | Vertical travel range: min 35 cm to min 185 cm;
| 5.7.3 | Grid Pb 10/80; f0 = 115 cm and 150 cm/180 cm, stationary grid;
| 5.7.4 | Front-cover-to-detector distance: < 40 mm;
| 5.7.5 | X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
| 5.7.6 | Automatic exposure control: Standard.
| 5.7.7 | Cassette size: accept all sizes from 5”x7” to 14”x17”.
| 5.7.8 | Movement arrested by electromagnetic brakes.
| 5.8 | Patient Table:
| 5.8.1 | Table height: min 72 cm;
| 5.8.2 | Tabletop width: min 80 cm;
| 5.8.3 | Tabletop length: min 220 cm;
| 5.8.4 | X-ray absorption: ≤ 1.2 mm Al eq at 100 kV;
| 5.8.6 | Tabletop travel: mon two panes, ± 46 cm, longitudinal, ± 10 cm, transverse;
| 5.8.7 | Max. patient weight: min 200 kg;
| 5.8.8 | Grid Pb 10/80; f0 = 115 cm, stationary grid;
| 5.8.9 | Max. patient coverage: min 185 cm without patient repositioning;
| 5.8.10 | Detector cover range (edge-to-edge): ≥ 55 cm, movement arrested by electromagnetic brakes;
| 5.8.11 | Tabletop – detector distance: ≤ 70 mm;
| 5.9 | Workstation (Standard configuration):
| 5.9.1 | The imaging station is specially designed for diagnostic radiology and delivers excellent image quality with a fast and seamless workflow;
| 5.9.2 | High-resolution digital radiography system with DICOM network connection for image processing and image display on a preview monitor, min 2 MP;
| 5.9.3 | High performance PC included;
| 5.9.4 | DR printer
| 6 | Accessories, spares and consumables
| 6.1 | All standard accessories, consumables and parts required to operate the equipment.
| 7 | Power supply:
| 7.1 | Power supply: 415 ± 5%V (3 Phase), 50Hz
| 8 | Environmental factors
| 8.1 | The unit shall be capable of being stored continuously in ambient temperature of 0 -55deg C and relative humidity of 15-95%
| 8.2 | The unit shall be capable of operating in ambient temperature of 20-40 deg C and relative humidity of 80%
| 9 | Standards & Safety Requirements
| 9.1 | Should be CE or FDA approved product certificate.
| 9.2 | Must submit ISO13485
| 10 | Training
| 10.1 | Must provide user & service training.
| 11 | Warranty
| 11.1 | Comprehensive warranty for 2 years.
| 12 | Documentation
| 12.1 | User (Operating) manual in English Should provide 2 sets (hardcopy and soft-copy)
| 12.2 | Service (Technical / Maintenance) manual in English Should provide 2 sets (hardcopy and soft-copy)
| 12.3 | List of important spare parts, accessories and consumables with their part numbers and costing.